## Sheet1
| PRODUCT CODE | Product Code Name | Regulation Number | WEB_ADDRESS | RECALL_NUMBER | PRODUCT_DESCRIPTION | TRADE_NAME | RECALL_CLASS | CENTER_CLASSIFICATION_DT | POSTED_INTERNET_DT | TERMINATION_DT | FEI_NUMBER | FIRM_NAME | MANUFACTURER_RECALL_REASON
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=141378 | Z-0361-2016 | Model 5001102 MPS Delivery Set w/ arrest agent and additive cassettes, heat exchanger &10 ft. delivery tubing; Model 5001102-AS MPS Delivery Set with 6 ft. delivery tubing; Model 7001102 MPS Low Volume Delivery Set for use with cardiopulmonary bypass equipment. | MPS Delivery Set | 2 | 12/02/2015 | 12/02/2015 | 11/17/2016 | 3001665800 | Quest Medical, Inc. | The products have been found to intermittently exhibit a seal failure during use.
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=98620 | Z-0156-2015 | Terumo Custom Cardioplegia Delivery Set - Heat Exchanger Cardiopulmonary Bypass Item Numbers: 5852, 16010 Product Usage: This is a Cardioplegia Delivery System that is designed to delivery blood/solution to the patient undergoing cardiopulmonary bypass surgery. The set is designed to deliver blood/solutions at a 4:1 ratio. In includes a Conducer Heat Exchanger and a Bubble Trap module. Additionally, the set is constructed with a female luer connector that is the subject of the correction/removal activity | Modified MP-4Cardioplegia Delivery System with Conducer | 2 | 10/31/2014 | 10/31/2014 | 11/06/2014 | 1000149028 | Terumo Cardiovascular Systems Corp | Defect in the luer connector can compromise the seal on the female side of the connector and the compromised seal can introduce air in the table line
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=126931 | Z-0161-2015 | Terumo Custom Cardioplegia Delivery Sets Catalog Code: 204880 Product Usage: This is a replacement line to multiple Cardioplegia Delivery Sets-and includes the female luer connector" that is the subject of this correction/removal activity | MiniPrime Cardioplegia Set with Conducer | 2 | 10/31/2014 | 10/31/2014 | 11/06/2014 | 1000149028 | Terumo Cardiovascular Systems Corp | Defect in the luer connector can compromise the seal on the female side of the connector and the compromised seal can introduce air in the table line
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=98879 | Z-0158-2015 | Terumo Cardioplegia Delivery Set- Item Number: 16130 Product Usage: The Cardioplegia Delivery Sets are indicated for delivery of cardioplegic solutions to the heart during cardiopulmonary bypass surgery. | Clear Cardioplegia Delivery Sets | 2 | 10/31/2014 | 10/31/2014 | 11/06/2014 | 1000149028 | Terumo Cardiovascular Systems Corp | Defect in the luer connector can compromise the seal on the female side of the connector. and the compromised seal can introduce air in the table line
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93576 | Z-2614-2010 | Cardiovascular Procedure Kit X-Coated Pump and Table Pack, Item #71552 |  | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93578 | Z-2616-2010 | Cardiovascular Procedure Kit Custom 4:1 Myocardial Set, (MP4 w/ Coil Component Utilized), Item # 63485 | Cardiovascular Procedure Kit***Terumo Cardiovascular Systems Corporation | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93577 | Z-2615-2010 | Cardiovascular Procedure Kit, (MP4 w/ Conducer Component Utilized), labeled in part ***RWJ Hospital***, Item # 71305-03 | Cardiovascular Procedure Kit***Terumo Cardiovascular Systems Corporation | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93573 | Z-2612-2010 | Sarns Cardioplegia Set with w/ PVC Coil from MP4, Item # 15501 | MYO Cardioplegia Delivery Sets | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93584 | Z-2617-2010 | Cardiovascular Procedure Kit X-Coated H/L pack with Cardioplegia, Item # 70067-03 | Cardiovascular Procedure Kit***Terumo Cardiovascular Systems Corporation | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93574 | Z-2613-2010 | Sarns Cardioplegia Set CPLG BLD MP4 4:1 No recirculation, Item # 16010 | MYO Cardioplegia Delivery Sets | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93572 | Z-2611-2010 | Sarns Cardioplegia Set with Conducer Heat Exchanger Cplg/MP4 4:1, Item # 165720 | MYO Cardioplegia Delivery Sets | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93571 | Z-2610-2010 | Sarns Cardioplegia Set Conducer Heat Exchanger Cplg/MP4 4:1 Bridge, Item# 6375 | MYO Cardioplegia Delivery Sets | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92559 | Z-2163-2010 | Terumo Cardiovascular Procedure Kit Custom CONDUCER/MP4 8:1 NO RECIRC. DOUBLE SPIKE P/N: 71032 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92558 | Z-2162-2010 | Terumo Cardiovascular Procedure Kit Custom X-COATED HEART - LUNG PACK P/N: 70608-01 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92498 | Z-2157-2010 | Terumo Cardiovascular Procedure Kit Custom X-COATED PERFUSION PACK P/N 71934 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92482 | Z-2150-2010 | Terumo Cardiovascular Procedure Kit Custom 1:1 CONDUCER/MP4 CRYSTALLOID CARDIOPLEGIA P/N 65246 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92504 | Z-2159-2010 | Terumo Cardiovascular Procedure Kit Custom X-COATED MASTER PACK P/N 72790 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92497 | Z-2156-2010 | Terumo Cardiovascular Procedure Kit Custom X-COATED HEART/LUNG FRAMEPACK P/N 71766PP | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92487 | Z-2151-2010 | Terumo Cardiovascular Procedure Kit Custom MP4, 4:1, PRELIEF, W/RECIRC, COND / PUMP & TABLE PACK P/N 65503/70881-02 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92507 | Z-2161-2010 | Terumo Cardiovascular Procedure Kit Custom ADULT PACK P/N: 131351-02 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92506 | Z-2160-2010 | Terumo Cardiovascular Procedure Kit Custom X-COATED ADULT PACK P/N 72846 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92494 | Z-2154-2010 | Terumo Cardiovascular Procedure Kit Custom ADULT PACK P/N 71546-01 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92492 | Z-2153-2010 | Terumo Cardiovascular Procedure Kit Custom X-COATED ADULT PRESTIGE PACK P/N 71544-01 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92490 | Z-2152-2010 | Terumo Cardiovascular Procedure Kit Custom 8:1 CARDIOPLEGIA SET P/N 70608-01 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92501 | Z-2158-2010 | Terumo Cardiovascular Procedure Kit Custom X-COATED PUMP AND TABLE PACK W/8:1 CARDIOPLEGIA P/N 72642 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92474 | Z-2149-2010 | Terumo Cardiovascular Procedure Kit Custom Open Perfusion Pack P/N 64937 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92495 | Z-2155-2010 | Terumo Cardiovascular Procedure Kit Custom PERFUSION TUBING PACK P/N 71733 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=200678 | Z-1641-2023 | QUADROX-i Pediatric, Oxygenator used during cardiopulmonary bypass Model Item No. HMO 30000-USA 701070384 HMO 31000-USA 701070388 | QUADROX-i Pediatric | 2 | 05/26/2023 | 05/26/2023 |  | 3008355164 | Maquet Medical Systems USA | 1. An accessory may be improperly placed during packaging, causing creases and damage to the accessory's sterile pouch. 2. Product packaging may exhibit holes, cracks, dents, and crushed areas. Both issues may compromise the sterile barrier of the packaging, and exposure to non-sterile product may result in inflammation, infection, sepsis, and ischemia.
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180728 | Z-2107-2020 | Ref 801168, MPS Top Cover Replacement Kit, 1 UNITS, UDI: (01)2 0634624 81168 7 (11)190916 (17)290913 (10)059135 (The device is a component of the MPS Console used as replacements for worn or damaged MPS Console top covers) | MPS2 Myocardial Protection System Console | 2 | 05/22/2020 |  | 04/17/2024 | 3001665800 | Quest Medical, Inc. | Additive and Arrest Agent Labels on the MPS Top Cover Replacement Kit, a replacement for worn or damaged MPS Console Top Covers for cardiopulmonary bypass heat exchanger, used to deliver whole blood to the heart during open heart surgery, were swapped and incorrectly placed. This could lead to over delivery of potassium by the user requiring medical intervention.
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=95458 | Z-1924-2011 | Quest MPS and MPS 2 Consoles, Quest Medical, Inc., An Atrion company. Single software controlled device that incorporates a pump, a heater/cooler, temperature monitoring, pressure monitoring, a heat exchanger, an arrest agent pump and an additive pump. The MPS console is used with the MPS sterile disposables with a built in heat exchanger. The device is used with on-pump or off-pump cardiac surgery. | MPS and MPS 2 Consoles | 3 | 04/08/2011 | 04/08/2011 | 04/26/2011 | 3001665800 | Quest Medical, Inc | User update regarding installation of external heater-Cooler units with MPS and MPS2 consoles.
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=25116 | Z-0298-03 | Heat Exchanger Coil (code Number 0006-0011) used in the ''Custom Cardiovascular Procedure Kits'' . | Terumo Cardiovacular Systems Heat Exchanger Coil | 2 | 04/02/2003 | 04/03/2003 | 08/06/2015 | 1000149028 | Terumo Cardiovascular Systems Corp | Heat exchanger may contain loose particulate matter
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=26348 | Z-0664-03 | Sarns MP4 Cardioplegia Sets; catalog number 15501. | Sarns MP4 Cardioplegia Sets | 2 | 03/19/2003 | 03/20/2003 | 06/06/2003 | 1828100 | Terumo Cardiovascular Systems Corp | There is a potential for delivery of a lower than intended amount of cardioplegia solution to the heart during surgery.
2 | BSK | CUFF, TRACHEAL TUBE, INFLATABLE | 868.5750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=125652 | Z-1165-2014 | MAQUET PLEGIOX Cardioplegia Heat Exchanger The PLEGIOX Cardioplegia Heat Exchanger is used as part of a cardioplegia system to set and maintain temperature for given flow rates and within the given temperature range of blood cardioplegic and crystalloid cardioplegic solutions during extracorporeal circulation. The product is designed for single use only, for an application period of no longer than 6 hours. | MAQUET | 2 | 03/10/2014 | 03/10/2014 | 04/24/2015 | 3008355164 | Maquet Cardiovascular Us Sales, Llc | The MAQUET post-marketing vigilance program has identified a discrepancy that the desired cardioplegia temperature may not be obtained when the PLEGIOX Cardioplegia Heat Exchanger (PLEGIOX) is used in conjunction with a heater/cooler unit.
2
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=95923 | Z-0813-2011 | Blease Sirius Anaesthesia System is a anaesthesia Gas Machine. The system has the capability to mount two anaesthetic vapourisers on the back bar via a demountable system. The delivery of anaesthetic agent is achieved by the rotation of a control knob which actuates a feature designed "shot bold" mechanism. The purpose of which is to prevent two vapourisers being turned on simultaneously.. The product is labeled in part: "BleaseSirius Anesthesia System...SPACELABS Healthcare...Blease Medical Equipment Limited Beech House, Chiltern Court, Asheridge Road, Chesham, Bucks HP5 2PX, England...Spacelabs Healthcare Company Headquarters 5150 220th Ave. SE, Issaquah, WA 98029 PO Box 7018, Issaquah, WA 98027-7018". | Blease Sirius Anaesthesia Systems | 2 | 12/27/2010 | 12/27/2010 | 12/28/2010 | 3010157426 | Spacelabs Healthcare, Llc | Users may experience difficulty turning the control knob to an active vapour delivery setting, or may not be able to turn on the Left Hand mounted vapouriser on the Blease Sirius Anaesthesia Systems.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=114600 | Z-0565-2013 | BleaseFocus Anesthesia Workstation and 700/900 Series Ventilators are anesthesia gas machine. In most cases, the customer ordered both a BleaseFocus anesthesia workstation and a 700/900 ventilator. In some cases, the customer may order a 700/900 stand-alone ventilator would could be used with a BleaseFocus anesthesia workstation at a later date. | BleaseFocus Anesthesia Workstation and 700/900 Series Ventilators | 2 | 12/19/2012 | 12/19/2012 | 07/11/2013 | 3002807992 | Del Mar Reynolds Medical, Ltd. | The BleaseFocus Anesthesia Workstation and 700/900 Series Ventilators were shipped with manuals that describe a secondary O2 Supply Failure alarm which has not been incorporated into the system. The manuals were prematurely updated for a feature that is not yet available.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=217545 | Z-0914-2026 | GE Healthcare Carestation anesthesia system, product number and REF numbers: Carestation 750 A1, REF 1012-9750-000; Carestation 750 A2, REF 1012-9750-002; Carestation 750c A1, REF 1012-9755-000; Products for which affected PMB can be used as a spare part: Carestation 750c A2, REF 1012-9755-002 | GE Healthcare Carestation anesthesia system | 1 | 12/18/2025 | 12/18/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Potential unexpected shutdown of Carestation 600 and 700 Series Anesthesia Systems containing certain power management boards, if the AC mains power is unplugged or in the event of an AC mains power failure. Anesthesia systems only operate on battery power in a rare event that AC mains power is lost and there is no continuous backup emergency power. If AC power is interrupted, the system will not automatically switch over to the battery supply mode and will reboot. If this issue occurs, a temporary disruption of mechanical ventilation, manual ventilation, and volatile agent delivery may occur. Following the reboot, the system will not return to the previous ventilation settings.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=217546 | Z-0915-2026 | GE Healthcare Carestation anesthesia system Power Management Board (PMB) Field Replaceable Unit (FRU), 2076139-001-S | GE Healthcare Carestation anesthesia system | 1 | 12/18/2025 | 12/18/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Potential unexpected shutdown of Carestation 600 and 700 Series Anesthesia Systems containing certain power management boards, if the AC mains power is unplugged or in the event of an AC mains power failure. Anesthesia systems only operate on battery power in a rare event that AC mains power is lost and there is no continuous backup emergency power. If AC power is interrupted, the system will not automatically switch over to the battery supply mode and will reboot. If this issue occurs, a temporary disruption of mechanical ventilation, manual ventilation, and volatile agent delivery may occur. Following the reboot, the system will not return to the previous ventilation settings.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=217544 | Z-0913-2026 | GE Healthcare Carestation anesthesia system, product number and REF numbers: Carestation 610 A1, REF 1012-9620-222; Carestation 620 A1 REF 1012-9620-000; Carestation 620 A1, REF 1012-9620-200; Carestation 620 A1, REF 1012-9620-202; Carestation 620 SE A1, REF 1012-9620-212; Carestation 630 A1, REF 1012-9650-222; Carestation 650 A1, REF 1012-9650-000; Carestation 650 A1, REF 1012-9650-200; Carestation 650 A1, REF 1012-9650-202; Carestation 650 SE A1, REF 1012-9650-212; Carestation 650c A1, REF 1012-9655-202; Anesthesia Gas Machine Products for which affected PMB can be used as a spare part: Carestation 620 A2, REF 1012-9620-002 00840682124546 Carestation 620 SE A2, REF 1012-9620-012 00195278569677 Carestation 650 A2, REF 1012-9650-002 00840682124560 Carestation 650 SE A2, REF 1012-9650-012 00195278569684 Carestation 650c A1, REF 1012-9655-000 00840682103954 Carestation 650c A2, REF 1012-9655-002 00840682124539 Carestation 650c A1, REF 1012-9655-200 00195278439543 | GE Healthcare Carestation anesthesia system | 1 | 12/18/2025 | 12/18/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Potential unexpected shutdown of Carestation 600 and 700 Series Anesthesia Systems containing certain power management boards, if the AC mains power is unplugged or in the event of an AC mains power failure. Anesthesia systems only operate on battery power in a rare event that AC mains power is lost and there is no continuous backup emergency power. If AC power is interrupted, the system will not automatically switch over to the battery supply mode and will reboot. If this issue occurs, a temporary disruption of mechanical ventilation, manual ventilation, and volatile agent delivery may occur. Following the reboot, the system will not return to the previous ventilation settings.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211121 | Z-0667-2025 | Brand Name: Atlan Product Name: Atlan A300 Model/Catalog Number: 8211300 The Atlan is an anesthesia workstation intended for use in anesthetizing adults, pediatrics, and neonates. | Atlan | 1 | 12/18/2024 | 12/18/2024 |  | 2517967 | Draeger Medical, Inc. | The Atlan anesthesia workstation either indicated a failure of the piston ventilator before use or suffered a failure of the mechanical ventilation during use.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211123 | Z-0668-2025 | Brand Name: Atlan Product Name: Atlan A300XL Model/Catalog Number: 8621400 The Atlan is an anesthesia workstation intended for use in anesthetizing adults, pediatrics, and neonates. | Atlan | 1 | 12/18/2024 | 12/18/2024 |  | 2517967 | Draeger Medical, Inc. | The Atlan anesthesia workstation either indicated a failure of the piston ventilator before use or suffered a failure of the mechanical ventilation during use.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=48422 | Z-0291-2007 | Datex Ohmeda Compact Absorber, Disposable, REF 427002100, white to violet and REF 427002000, Pink to white. Used with GE Healthcare ADU anesthesia systems. GE Healthcare Finland Oy, Helsinki, Finland +358 10 39411. | Compact Absorbers | 2 | 12/18/2006 | 12/19/2006 | 12/17/2011 | 2183066 | Datex - Ohmeda, Inc | Certain Compact Absorbers may have an increased resistance to gas flow due to an improperly manufactured foam filter. The increased resistance can cause an elevated pressure at the ventilator end of the inspiratory circuit, but the pressure at the patient may be reduced. This could result in patient hypoventilation and hypoxia.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175185 | Z-0116-2020 | Aespire 7100, Model Numbers: 1. 1009-9000-000-009469 2. 1009-9000-000 3. 1009-9000-000-013913 4. 1009-9000-000-001366 5. 1009-9000-000-009857 6. 1009-9003-000 7. 1009-9000-000-008120 8. 1009-9000-000-009600 9. 1009-9000-000-014290 10. 1009-9000-000-010585 11. 1009-9000-000-004993 12. 1009-9011-000-004795 13. 1009-9000-000-013407 14. 1009-9000-000-004604 15. 1009-9000-000-017674 16. 1009-9000-000-008763 17. 1009-9000-000-008306 18. 1009-9000-000-003277 19. 1009-9000-000-004166 20. 1009-9000-000-008560 21. 1009-9000-000-000299 22. 1009-9000-000-011444 23. 1009-9000-000-003529 24. 1009-9000-000-005666 25. 1009-9000-000-019821 26. 1009-9011-000 27. 1009-9000-000-017345 28. 1009-9000-000-009893 29. 1009-9000-000-007258 30. 1009-9000-000-004959 31. 1009-9000-000-010185 32. 1009-9000-000-016672 33. 1009-9011-000-017769 34. 1009-9000-000-014329 35. 1009-9011-000-020685 36. 1009-9000-000-026022 37. 1009-9011-000-002912 38. 1009-9000-000-014723 39. 1009-9000-000-013385 40. 1009-9000-000-006725 41. 1009-9000-000-005163 42. 1009-9000-000-004022 43. 1009-9000-000-017939 44. 1009-9000-000-000300 45. 1009-9000-000-018855 46. 1009-9000-000-016724 47. 1009-9000-000-010268 48. 1009-9000-000-011595 49. 1009-9000-000-019277 50. 1009-9011-000-023222 51. 1009-9011-000-015929 52. 1009-9000-000-017824 53. 1009-9000-000-008270 54. 1009-9000-000-010416 55. 1009-9000-000-019089 56. 1009-9000-000-013395 57. 1009-9011-000-021989 58. 1009-9000-000-000517 59. 1009-9000-000-002407 60. 1009-9000-000-008415 61. 1009-9000-000-009597 62. 1009-9000-000-010178 63. 1009-9011-000-010400 64. 1009-9011-000-011366 65. 1009-9000-000-012360 66. 1009-9101-000 67. 1009-9000-000-013787 68. 1009-9000-000-006266 69. 1009-9000-000-014915 70. 1009-9000-000-000290 71. 1009-9000-000-013507 72. 1009-9000-000-011600 73. 1009-9000-000-010595 74. 1009-9000-000-009543 75. 1009-9000-000-011111 76. 1009-9000-000-013378 77. 1009-9000-000-000526 78. 1009-9000-000-010536 79. 1009-9000-000-010116 80. 1009-9000-000-005939 81. 1009-9000-000-008216 82. 1009-9000-000-006786 83. 1009-9000-000-012155 84. 1009-9000-000-012379 85. 1009-9000-000-014000 86. 1009-9011-000-022466 87. 1009-9000-000-009912 88. 1009-9000-000-006906 89. 1009-9000-000-004380 90. 1009-9000-000-010245 91. 1009-9000-000-010413 92. 1009-9000-000-013616 93. 1009-9000-000-009428 94. 1009-9000-000-010597 95. 1009-9000-000-020606 96. 1009-9000-000-010059 97. 1009-9000-000-014122 98. 1009-9000-000-016795 99. 1009-9000-000-012149 100. 1009-9000-000-005760 101. 1009-9000-000-009836 102. 1009-9000-000-016818 103. 1009-9000-000-017699 104. 1009-9000-000-004020 105. 1009-9000-000-013953 106. 1009-9000-000-009583 107. 1009-9000-000-007938 108. 1009-9000-000-007336 109. 1009-9000-000-010435 110. 1009-9011-000-016719 111. 1009-9000-000-024098 112. 1009-9011-000-004634 113. 1009-9011-000-016794 114. 1009-9011-000-005431 115. 1009-9000-000-000598 116. 1009-9011-000-005442 117. 1009-9000-000-022792 118. 1009-9000-000-003950 119. 1009-9000-000-013921 120. 1009-9000-000-011116 121. 1009-9000-000-017203 122. 1009-9000-000-003335 123. 1009-9000-000-008264 124. 1009-9000-000-009868 125. 1009-9011-000-015245 truncated text | Aespire 7100 | 2 | 12/17/2019 |  | 04/13/2026 | 2126677 | GE Healthcare, LLC | Certain Aespire and Aestiva Anesthesia Systems were noted to have a vulnerability to a cyber-attack when connected to the hospital network. An insufficiently secured terminal server may provide an opportunity for a malicious actor that has already penetrated the hospital network to send fraudulent flow sensor correction parameters.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175186 | Z-0117-2020 | Aestiva 7100, Model Numbers: 1. 1006-9320-000 2. 1006-9321-000 3. 1006-9320-000-001101 4. 1006-9021-000 5. 1006-9012-000 6. 1006-9111-000 7. 1006-9321-000-007639 8. 1006-9321-000-000023 9. 1006-9321-000-015180 10. 1006-9320-000-010369 11. 1006-9321-000-019486 12. 1006-9321-000-007601 13. 1006-9321-000-009266 14. 1006-9321-000-008526 15. 1006-9321-000-116507 16. 1006-9321-000-009244 17. 1006-9321-000-012769 18. 1006-9320-000-004497 19. 1006-9320-000-009697 20. 1006-9320-000-014484 21. 1006-9321-000-008072 22. 1006-9321-000-007618 23. 1006-9321-000-013784 24. 1006-9320-000-010963 25. 1006-9320-000-014485 26. 1006-9320-000-008839 27. 1006-9321-000-017374 28. 1006-9321-000-014553 29. 1006-9031-000 30. 1006-9321-000-000533 31. 1006-9321-000-007785 32. 1006-9320-000-016418 33. 1006-9320-000-013302 34. 1006-9320-000-020339 35. 1006-9320-000-002141 36. 1006-9320-000-009193 37. 1006-9320-000-014443 38. 1006-9320-000-008124 39. 1006-9321-000-009347 40. 1006-9320-000-009310 41. 1006-9320-000-004623 42. 1006-9321-000-020437 43. 1006-9321-000-025830 44. 1006-9321-000-004690 45. 1006-9321-000-011348 46. 1006-9114-000 47. 1006-9322-000 48. 1006-9321-000-006678 49. 1006-9321-000-043401 50. 1006-9322-000-043399 51. 1006-9320-000-018778 52. 1006-9320-000-010161 53. 1006-9320-000-009400 54. 1006-9320-000-007948 55. 1006-9321-000-001309 56. 1006-9321-000-008305 57. 1006-9321-000-027472 58. 1006-9321-000-008643 59. 1006-9320-000-008841 60. 1006-9320-000-010102 61. 1006-9322-000-020407 62. 1006-9321-000-004831 63. 1006-9320-000-010684 64. 1006-9320-000-014654 65. 1006-9320-000-006392 66. 1006-9320-000-009056 67. 1006-9321-000-025850 68. 1006-9320-000-000473 69. 1006-9320-000-017639 70. 1006-9321-000-003110 71. 1006-9321-000-019653 72. 1006-9321-000-008529 73. 1006-9321-000-010029 74. 1006-9321-000-005790 75. 1006-9320-000-010177 76. 1006-9320-000-028131 77. 1006-9321-000-006318 78. 1006-9322-000-106884 79. 1006-9320-000-004596 80. 1006-9321-000-006645 81. 1006-9321-000-010584 82. 1006-9320-000-007779 83. 1006-9320-000-025679 84. 1006-9320-000-011509 85. 1006-9320-000-008650 86. 1006-9321-000-020045 87. 1006-9321-000-014996 88. 1006-9320-000-015449 89. 1006-9320-000-010685 90. 1006-9321-000-000266 91. 1006-9321-000-019680 92. 1006-9321-000-004778 93. 1006-9321-000-004424 94. 1006-9321-000-008747 95. 1006-9320-000-009249 96. 1006-9320-000-010834 97. 1006-9320-000-022596 98. 1006-9024-000 99. 1006-9026-000 100. 1006-9322-000-011934 101. 1006-9322-000-049412 102. 1006-9322-000-110551 103. 1006-9320-000-019424 104. 1006-9321-000-017187 105. 1006-9321-000-005997 106. 1006-9321-000-091883 107. 1006-9321-000-004524 108. 1006-9321-000-013425 109. 1006-9320-000-015413 110. 1006-9321-000-007378 111. 1006-9320-000-010661 112. 1006-9321-000-005998 113. 1006-9321-000-156642 114. 1006-9320-000-015195 115. 1006-9320-000-017086 116. 1006-9320-000-017852 117. 1006-9321-000-001974 118. 1006-9321-000-010236 119. 1006-9321-000-007641 120. 1006-9321-000-014028 121. 1006-9321-000-015552 122. 1006-9321-000-011524 123. 1006-9321-000-009297 124. 1006-9320-000-017427 125. 1006-9320-000-017498 126. 1006-9321-000-022470 127. 1006-9321-000-024751 128. 1006-9321-000-007642 129. 1006-9321-000-007636 130. 1006-9321-000-029968 131. 1006-9321-000-039161 132. 1006-9321-000-035978 133. 1006-9321-000-011702 134. 1006-9321-000-043612 models truncated | Aestiva 7100 | 2 | 12/17/2019 |  | 04/13/2026 | 2126677 | GE Healthcare, LLC | Certain Aespire and Aestiva Anesthesia Systems were noted to have a vulnerability to a cyber-attack when connected to the hospital network. An insufficiently secured terminal server may provide an opportunity for a malicious actor that has already penetrated the hospital network to send fraudulent flow sensor correction parameters.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=175183 | Z-0114-2020 | Aestiva MRI, Model Numbers: a) 1006-9310-000 b) 1006-9110-000 c) 1006-9023-000 d) 1006-9028-000 e) 1006-9310-000-305077 f) 1006-9310-000-015243 g) 1006-9310-000-017602 h) 1006-9310-000-103785 I) 1006-9310-000-025109 j) 1006-9310-000-009650 k) 1006-9310-000-015224 l) 1006-9310-000-031881 m) 1006-9310-000-031854 n) 1006-9310-000-026571 | Aestiva MRI | 2 | 12/17/2019 |  | 04/13/2026 | 2126677 | GE Healthcare, LLC | Certain Aespire and Aestiva Anesthesia Systems were noted to have a vulnerability to a cyber-attack when connected to the hospital network. An insufficiently secured terminal server may provide an opportunity for a malicious actor that has already penetrated the hospital network to send fraudulent flow sensor correction parameters.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=150848 | Z-0755-2017 | GE Healthcare, Avance, Avance CS2, Amingo. | GE Healthcare, Avance, Avance CS2, Amingo. | 2 | 12/14/2016 |  | 09/17/2018 | 2126677 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue where certain Avance CS2, Avance and Amingo anesthesia devices can transition to a System Malfunction state if the lower storage drawer containing the optional large tray insert accessory closed with an abnormally high amount of force.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=141791 | Z-0360-2016 | Perseus A500 Anesthesia Workstation, MK06000. Intended for use in anesthetizing adults, children, and neonates and can be used for automatic and manual ventilation, pressure-supported spontaneous breathing, and spontaneous breathing. | Perseus A500 Anesthesia Workstation | 1 | 12/11/2015 | 12/12/2015 | 04/05/2017 | 2517967 | Draeger Medical, Inc. | There is a potential problem with the device main power switch of the Perseus A500 Anesthesia Workstation. The device power switch may spontaneously fail during use, causing one of the following to occur: 1. The workstation has a fresh gas delivery failure and a ventilator failure with respective alarms. 2. The workstation shuts down completely with an audible alarm.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=196696 | Z-0422-2023 | Flow-c Anesthesia System (Product Code 6887700) Flow-e Anesthesia System (Product Code 6887900) | Flow-c, Flow-e | 2 | 12/07/2022 | 12/07/2022 |  | 3013876692 | Getinge Usa Sales Inc | The power backup battery, under certain conditions, may exhibit reduced battery run-time while in use, and there is a potential that the battery may fail to provide the expected run time in the event of a mains power failure. This could require the clinician to use manual emergency ventilation of another ventilator is not available and the backup battery is exhausted.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=74064 | Z-0319-2009 | AS3000 Anesthesia Delivery System. Datascope Patient Monitoring; A Mindray Global Company. The AS3000 is a device used to administer to a patient, continuously or intermittently, a general inhalation anesthetic and to maintain a patient's ventilation. | NAVIGATOR ANESTHESIA DELIVERY SYSTEM | 2 | 12/05/2008 | 12/05/2008 | 12/08/2008 | 2221819 | Mindray DS USA, Inc., dba Datascope Patient Monitoring | Two issues have been identified with the AS3000 unit. 1. The threadlocker of the caster (wheels) may not have been utilized, which may allow the caster to loosen and possibly separate from the unit. 2. The use of select brands of pre-pack absorber in the absorber canister of the AS3000 has been associated, in some cases, with gas leakage around the pre-pack, rather than through the absorber mate
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=216622 | Z-0720-2026 | Sevoflurane (SEVO) V90 Electronic Vaporizer with the A9 Anesthesia System. Labeled as the following: 1. Sev E_Vap. (Safety Filling adapter), Part Number: 115-066756-00. 2. Exchange Sev E_Vap. (Safety Filling adapter), Part Number: E115-066756-00. Product Description: The Anesthesia System is a device used to administer to a patient, continuously or intermittently, a general inhalation anesthetic and to maintain a patient's ventilation. The A9 Anesthesia System utilizes V90 electronic vaporizers for anesthetic agent delivery. | Sevoflurane (SEVO) V90 Electronic Vaporizer | 2 | 11/26/2025 | 11/26/2025 |  | 2221819 | Mindray DS USA, Inc. dba Mindray North America | Potential for anesthesia leakage.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=216623 | Z-0721-2026 | Isoflurane (ISO) V90 Electronic Vaporizer with the A9 Anesthesia System. Labeled as the following: 1. Iso E_Vap. (Safety Filling adapter), Part Number 115-066758-00. 2. Exchange Iso E_Vap. (Safety Filling adapter), Part Number E115-066758-00. Product Description: The Anesthesia System is a device used to administer to a patient, continuously or intermittently, a general inhalation anesthetic and to maintain a patient's ventilation. The A9 Anesthesia System utilizes V90 electronic vaporizers for anesthetic agent delivery. | Isoflurane (ISO) V90 Electronic Vaporizer | 2 | 11/26/2025 | 11/26/2025 |  | 2221819 | Mindray DS USA, Inc. dba Mindray North America | Potential for anesthesia leakage.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=159202 | Z-0072-2018 | Arkon Anesthesia Delivery System, Model 99999, and software version 2.61 and upgrade kit 050-9043-00 Rev G. The Spacelabs Arkon Anesthesia Workstation is intended for use in the hospital environment and operating room. It may be used for the delivery of oxygen, air and nitrous oxide in a controlled manner to various patient breathing circuits with or without the use of mechanical ventilator, and may be used for the delivery of anesthetic vapor by use of a dismountable vaporizer. The device is intended for use only by a suitably qualified physician | Arkon Anesthesia Delivery System and Upgrade Kit 050-9043-00 | 1 | 11/21/2017 | 11/21/2017 | 03/12/2018 | 3002807992 | Spacelabs Healthcare, Ltd. | Arkon Anesthesia Workstation, with software version 2.61, experienced failure in mechanical ventilation, oxygen and anesthetic gas delivery, with concurrent failure of the display unit that resulted in a blank screen without audible or visible alarms.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=123126 | Z-0352-2014 | CAS I/II Absorbers that are parts of the BleaseSirius Anesthesia Workstation, the BleaseFocus Anesthesia Workstation with the following part numbers: 12200900; 12200901; 12200902; 12200903; 14200100, and service kits part number 050-0659-00 and 050-0901-00. BleaseFocus Anesthesia Workstation: The Blease Frontline Plus Range, Anesthesia Machines are intended for use in the hospital environment and locations not requiring portability. It may be used for the delivery of oxygen, air and nitrous oxide in a controlled manner to various patient breathing circuits with or without the use of a mechanical ventilator, and may be used for the delivery of anesthetic vapor by use of a dismountable vaporizer. This device is intended for use only by a suitably qualified physician. BleaseSirius Anesthesia Workstation: The Spacelabs BleaseSirius Anesthesia Workstation is intended for use in the hospital environment and operating room. It may be used for the delivery of oxygen, air and nitrous oxide in a controlled manner to various patient breathing circuits with or without the use of mechanical ventilator, and may be used for the delivery of anesthetic vapor by use of a dismountable vaporizer. The device is intended for use only by a suitably qualified physician. | BleaseSirius Anesthesia Workstation, BleaseFocus Anesthesia Workstation | 1 | 11/20/2013 | 11/20/2013 | 04/24/2014 | 3002807992 | Del Mar Reynolds Medical, Ltd. | Spacelabs Healthcare Ltd. of United Kingdom recalls , and Service Kits, Part Number 050-0659-00 and 050-0901-00. The Bag to-Vent switch in CAS I/II Absorbers may fail due to loose fastening hardware in an absorber.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=113481 | Z-0290-2013 | Mindray A3 Anesthesia Delivery System Manufactured in China for Mindray DS USA Inc. 800 MacArthur Boulevard, Mahwah, NJ 07430 This is a device used to administer to a patient continuously or intermittently a general inhalation anesthetic and to maintain a patient's ventilation. | A3/A5 Anesthesia Delivery System | 1 | 11/16/2012 | 11/16/2012 | 06/04/2013 | 2221819 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified the potential for a system leak resulting an issue related to the carbon dioxide absorber canister gasket on A3/A5 Anesthesia Delivery Systems. A potential system leak resulting from the canister gasket will likely present itself during the Automatic Circuit Leak and Compliance Test performed at startup and the Manual Leak Test recommended to be performed before each case.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=113482 | Z-0291-2013 | Mindray A5 Anesthesia Delivery System Manufactured in China for Mindray DS USA Inc. 800 MacArthur Boulevard, Mahwah, NJ 07430 This is a device used to administer to a patient continuously or intermittently a general inhalation anesthetic and to maintain a patient's ventilation. | A3/A5 Anesthesia Delivery System | 1 | 11/16/2012 | 11/16/2012 | 06/04/2013 | 2221819 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified the potential for a system leak resulting an issue related to the carbon dioxide absorber canister gasket on A3/A5 Anesthesia Delivery Systems. A potential system leak resulting from the canister gasket will likely present itself during the Automatic Circuit Leak and Compliance Test performed at startup and the Manual Leak Test recommended to be performed before each case.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210599 | Z-0304-2025 | Brand Name: Atlan Product Name: Atlan A350 Model/Catalog Number: 8211500 Software Version: Not applicable Component: Not applicable | Atlan | 1 | 11/13/2024 | 11/13/2024 |  | 2517967 | Draeger Medical, Inc. | The Atlan anesthesia workstation either indicated a failure of the piston ventilator before use or suffered a failure of the mechanical ventilation during use.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210600 | Z-0305-2025 | Brand Name: Atlan Product Name: Atlan A350XL Model/Catalog Number: 8621600 Software Version: Not Applicable Component: Not applicable | Atlan | 1 | 11/13/2024 | 11/13/2024 |  | 2517967 | Draeger Medical, Inc. | The Atlan anesthesia workstation either indicated a failure of the piston ventilator before use or suffered a failure of the mechanical ventilation during use.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=35610 | Z-0192-05 | APL Valve used on some Fabius GS (Catalog number 4117110), Fabius Tiro (Catalog number 4118350) and Narkomed 6000 (Catalog number 4118070) series Anesthesia Machines. | APL Valve used on some Fabius GS, Fabius Tiro and Narkomed 6000 series Anesthesia Machines | 2 | 11/10/2004 | 11/11/2004 | 06/19/2007 | 2517967 | Draeger Medical, Inc. | rotating knob of the rotary style APL valve can become separated from the assembly.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=168354 | Z-0413-2019 | Touchscreen display kit. Sold under the following product names: TOUCHSCREEN DISPLAY KIT 11.X ENGLISH FOR USA, TOUCHSCREEN DISPLAY KIT 11.X ENGLISH, TOUCHSCREEN DISPLAY KIT 11.X GERMAN, TOUCHSCREEN DISPLAY KIT 11.X DUTCH, TOUCHSCREEN DISPLAY KIT 11.X FRENCH, TOUCHSCREEN DISPLAY KIT 11.X DANISH, TOUCHSCREEN DISPLAY KIT 11.X ESTONIAN, TOUCHSCREEN DISPLAY KIT 11.X ITALIAN, TOUCHSCREEN DISPLAY KIT 11.X NORWEGIAN. Product Usage: The GE Datex-Ohmeda Aisys Anesthesia System is intended to provide general inhalation anesthesia and ventilatory support to a wide range of patients (neonatal, pediatric, adult). The device is intended for volume or pressure control ventilation. The Aisys is not suitable for use in a MRI environment. | Aisys Anesthesia System | 2 | 11/09/2018 | 11/09/2018 | 04/12/2021 | 2126677 | GE Healthcare, LLC | Upgraded Aisys anesthesia devices and Aisys CS2 anesthesia devices with PSVPro Spontaneous Breathing Modes do not provide the PSVPro ventilation mode feature of transitioning to SIMV PCV backup ventilation mode. PSVPro Spontaneous Breathing Modes do not provide the feature of transitioning to SIMV PCV backup ventilation mode for a patient who stops spontaneous breathing if the PSVPro ventilation mode is in use prior to the clinician starting the Cycling Procedure.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167708 | Z-0411-2019 | Aisys Anesthesia System. Sold under the following product names: AISYS, AISYS 10.X TO 11.X UPGRADE KIT DUTCH, AISYS 10.X TO 11.X UPGRADE KIT ENGLISH, AISYS 10.X TO 11.X UPGRADE KIT ENGLISH FOR USA, AISYS 10.X TO 11.X UPGRADE KIT FRENCH, AISYS 10.X TO 11.X UPGRADE KIT GERMAN, AISYS 10.X TO 11.X UPGRADE KIT JAPANESE. | Aisys Anesthesia System | 2 | 11/09/2018 | 11/09/2018 | 04/12/2021 | 2126677 | GE Healthcare, LLC | Upgraded Aisys anesthesia devices and Aisys CS2 anesthesia devices with PSVPro Spontaneous Breathing Modes do not provide the PSVPro ventilation mode feature of transitioning to SIMV PCV backup ventilation mode. PSVPro Spontaneous Breathing Modes do not provide the feature of transitioning to SIMV PCV backup ventilation mode for a patient who stops spontaneous breathing if the PSVPro ventilation mode is in use prior to the clinician starting the Cycling Procedure.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=167710 | Z-0412-2019 | Aisys CS2 Anesthesia System. Sold under the following product names: AISYS, AISYS 10.X TO 11.X UPGRADE KIT DUTCH, AISYS 10.X TO 11.X UPGRADE KIT ENGLISH, AISYS 10.X TO 11.X UPGRADE KIT ENGLISH FOR USA, AISYS 10.X TO 11.X UPGRADE KIT FRENCH, AISYS 10.X TO 11.X UPGRADE KIT GERMAN, AISYS 10.X TO 11.X UPGRADE KIT JAPANESE. Product Usage: The GE Datex-Ohmeda Aisys Anesthesia System is intended to provide general inhalation anesthesia and ventilatory support to a wide range of patients (neonatal, pediatric, adult). The device is intended for volume or pressure control ventilation. The Aisys is not suitable for use in a MRI environment. | Aisys Anesthesia System | 2 | 11/09/2018 | 11/09/2018 | 04/12/2021 | 2126677 | GE Healthcare, LLC | Upgraded Aisys anesthesia devices and Aisys CS2 anesthesia devices with PSVPro Spontaneous Breathing Modes do not provide the PSVPro ventilation mode feature of transitioning to SIMV PCV backup ventilation mode. PSVPro Spontaneous Breathing Modes do not provide the feature of transitioning to SIMV PCV backup ventilation mode for a patient who stops spontaneous breathing if the PSVPro ventilation mode is in use prior to the clinician starting the Cycling Procedure.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=47879 | Z-0021-2007 | GE Healthcare-Datex-Ohmeda Aisys Anesthesia Delivery System with Aladin 2 cassettes, GE Healthcare, Madison, WI; Datex-Ohmeda, inc., a General Electric company | GE Healthcare- Datex-Ohmeda Aisys | 2 | 11/06/2006 | 11/07/2006 | 02/24/2008 | 2183066 | Datex - Ohmeda, Inc | Vaporizer shutdown with alarm-An intermittent leak in the Aladin 2 cassette/Aisys pneumatic interface (valves) may occur that will invoke an alarm state. When this occurs 'Vaporizer Failure' will appear on the Aisys main display, with an accompanying audible alrm. Delivery of anesthetic Agent is halted; gas flow and ventilation are not affected.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=183922 | Z-0311-2021 | Aestiva/5 MRI Anesthesia System - Product Usage: The Aestiva/5 MRI Anesthesia System provides the functional feature set offered by the conventional Aestiva/5 to the clinician with the added ability to be used in the MR environment. Among those standard Aestiva/5 features is the Datex-Ohmeda user interface, all the ventilation parameters of the SmartVent along with the Aestiva breathing circuit. The Aestiva/5 MRI Anesthesia System performed to specifications when tested directly next to 1.5 and 3.0 Tesla shielded MRI devices in a field strength that did not exceed 300 gauss. | Aestiva MRI | 2 | 10/30/2020 |  |  | 2126677 | GE Healthcare, LLC | Flow sensors could have damaged tubes in the form of small punctures or cuts.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184237 | Z-0317-2021 | Flow Sensor Service Part, Catalog 2096513-001-S | Flow Sensor | 2 | 10/30/2020 |  |  | 2126677 | GE Healthcare, LLC | Flow sensors could have damaged tubes in the form of small punctures or cuts.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184236 | Z-0316-2021 | Flow Sensor Service Part, Catalog 2087640-001-S | Flow Sensor | 2 | 10/30/2020 |  |  | 2126677 | GE Healthcare, LLC | Flow sensors could have damaged tubes in the form of small punctures or cuts.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184235 | Z-0315-2021 | Flow Sensor Service Part, Catalog 2089610-001-S | Flow Sensor | 2 | 10/30/2020 |  |  | 2126677 | GE Healthcare, LLC | Flow sensors could have damaged tubes in the form of small punctures or cuts.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=183932 | Z-0312-2021 | GE Datex-Ohmeda Avance CS2 Anesthesia System - Product Usage: intended to provide general inhalation anesthesia and ventilatory support to a wide range of patients (neonatal, pediatric, adult). The device is intended for volume or pressure control ventilation. | Avance CS2 | 2 | 10/30/2020 |  |  | 2126677 | GE Healthcare, LLC | Flow sensors could have damaged tubes in the form of small punctures or cuts.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=183935 | Z-0314-2021 | 9100c NXT system - Product Usage: Device is licensed and distributed outside of the US only. | 9100c NXT | 2 | 10/30/2020 |  |  | 2126677 | GE Healthcare, LLC | Flow sensors could have damaged tubes in the form of small punctures or cuts.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=183934 | Z-0313-2021 | GE Datex-Ohmeda Aisys CS2 Anesthesia System - Product Usage: is intended to provide general inhalation anesthesia and ventilatory support to a wide range of patients (neonatal, pediatric adult). The device is intended for volume or pressure control ventilation. | Aisys CS2 | 2 | 10/30/2020 |  |  | 2126677 | GE Healthcare, LLC | Flow sensors could have damaged tubes in the form of small punctures or cuts.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209208 | Z-0080-2025 | Avance CS2 and Avance CS2 Pro. These anesthesia gas machines are intended to provide general inhalation anesthesia and ventilatory support to patients and are to be used only by medical professionals trained and qualified in the administration of general anesthesia. | Avance CS2 and Avance CS2 Pro | 1 | 10/18/2024 | 10/18/2024 |  | 2112667 | Datex-Ohmeda, Inc. | Potential issue that may arise if the limb of a patient breathing circuit is mistakenly connected to the Auxiliary Common Gas Outlet (ACGO) port.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209205 | Z-0077-2025 | Aespire View. This anesthesia gas machine is intended to provide general inhalation anesthesia and ventilatory support to patients and are to be used only by medical professionals trained and qualified in the administration of general anesthesia. | Aespire View | 1 | 10/18/2024 | 10/18/2024 |  | 2112667 | Datex-Ohmeda, Inc. | Potential issue that may arise if the limb of a patient breathing circuit is mistakenly connected to the Auxiliary Common Gas Outlet (ACGO) port.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209202 | Z-0074-2025 | Aespire 7100, 7100 (Variant), and Aespire 7100 S/5 Protiva configurations. These anesthesia gas machines are intended to provide general inhalation anesthesia and ventilatory support to patients and are to be used only by medical professionals trained and qualified in the administration of general anesthesia. | Aespire 7100 | 1 | 10/18/2024 | 10/18/2024 |  | 2112667 | Datex-Ohmeda, Inc. | Potential issue that may arise if the limb of a patient breathing circuit is mistakenly connected to the Auxiliary Common Gas Outlet (ACGO) port.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=210152 | Z-0084-2025 | Auxiliary Common Gas Outlet (ACGO) conversion/install kits sold under product names CONV KIT SCGO TO ACGO and ACGO ASSY FIELD INSTL. These kits are not affected specifically, however are included due to being purchased with or later used to convert to ACGO. |  | 1 | 10/18/2024 | 10/18/2024 |  | 2112667 | Datex-Ohmeda, Inc. | Potential issue that may arise if the limb of a patient breathing circuit is mistakenly connected to the Auxiliary Common Gas Outlet (ACGO) port.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209210 | Z-0081-2025 | Aisys. This anesthesia gas machine is intended to provide general inhalation anesthesia and ventilatory support to patients and are to be used only by medical professionals trained and qualified in the administration of general anesthesia. | Aisys | 1 | 10/18/2024 | 10/18/2024 |  | 2112667 | Datex-Ohmeda, Inc. | Potential issue that may arise if the limb of a patient breathing circuit is mistakenly connected to the Auxiliary Common Gas Outlet (ACGO) port.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209203 | Z-0075-2025 | Aespire 100. This anesthesia gas machine is intended to provide general inhalation anesthesia and ventilatory support to patients and are to be used only by medical professionals trained and qualified in the administration of general anesthesia. | Aespire 100 | 1 | 10/18/2024 | 10/18/2024 |  | 2112667 | Datex-Ohmeda, Inc. | Potential issue that may arise if the limb of a patient breathing circuit is mistakenly connected to the Auxiliary Common Gas Outlet (ACGO) port.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209204 | Z-0076-2025 | Aespire 7900. This anesthesia gas machine is intended to provide general inhalation anesthesia and ventilatory support to patients and are to be used only by medical professionals trained and qualified in the administration of general anesthesia. | Aespire 7900 | 1 | 10/18/2024 | 10/18/2024 |  | 2112667 | Datex-Ohmeda, Inc. | Potential issue that may arise if the limb of a patient breathing circuit is mistakenly connected to the Auxiliary Common Gas Outlet (ACGO) port.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209206 | Z-0078-2025 | Avance. This anesthesia gas machine is intended to provide general inhalation anesthesia and ventilatory support to patients and are to be used only by medical professionals trained and qualified in the administration of general anesthesia. | Avance | 1 | 10/18/2024 | 10/18/2024 |  | 2112667 | Datex-Ohmeda, Inc. | Potential issue that may arise if the limb of a patient breathing circuit is mistakenly connected to the Auxiliary Common Gas Outlet (ACGO) port.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209211 | Z-0082-2025 | Aisys CS2. This anesthesia gas machine is intended to provide general inhalation anesthesia and ventilatory support to patients and are to be used only by medical professionals trained and qualified in the administration of general anesthesia. | Aisys CS2 | 1 | 10/18/2024 | 10/18/2024 |  | 2112667 | Datex-Ohmeda, Inc. | Potential issue that may arise if the limb of a patient breathing circuit is mistakenly connected to the Auxiliary Common Gas Outlet (ACGO) port.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=48711 | Z-0048-2007 | Fabius GS Anesthesia Machine. Catalog number 8604699 | Fabius GS Anesthesia Machine | 2 | 10/18/2006 | 10/19/2006 | 08/07/2007 | 2517967 | Draeger Medical, Inc. | One of four casters may break loose from chassis
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121843 | Z-0010-2014 | GE Healthcare, Avance, Avance CS2, Aisys, Aespire View, and Engstrom. Intended for volume or pressure control ventilation. | GE Healthcare, Avance, Avance CS2, Aisys, Aespire View, and Engstrom | 1 | 10/15/2013 | 10/15/2013 | 01/30/2015 | 2126677 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving unresponsive buttons on the display of the Engstrm ventilator and the Aespire View, Aisys, and Avance anesthesia machines. As a result of a manufacturing issue, the buttons on the left, right, and bottom keypads may not always detect button presses. This may result in the inability to access certain menu functions wh
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=122228 | Z-0009-2014 | GE Healthcare, Aisys, Amingo, Avance, and Avance CS2. Intended for volume or pressure control ventilation. | GE Healthcare, Aisys, Amingo, Avance, and Avance CS2 | 1 | 10/15/2013 | 10/15/2013 | 10/01/2015 | 2126677 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the Avance, Avance CS2, and Aisys anesthesia devices. While operating in Pressure Control Ventilation  Volume Guarantee (PCV-VG) mode, the affected products may produce an over delivery tidal volume when given a unique sequence of user inputs and a collapsed bellows (partially filled circuit). Exposure to excessiv
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=42272 | Z-0038-06 | Datex-Ohmeda S/5 Aespire Anesthesia System 7900 | Datex-Ohmeda S/5 Aespire Anesthesia System 7100 | 2 | 10/11/2005 | 10/12/2005 | 04/01/2007 | 2183066 | Datex-Ohmeda Inc | Two conditions may occur with the S/5 Aespire Anesthesia System. The internal auxiliary flowmeter tubing may become disconnected if the auxiliary oxygen supply tubing is occluded and the patient may not receive desired supplemental oxygen which could lead to serious injury if undetected. Also the O2 switch may leak or fail.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=42273 | Z-0039-06 | Datex-Ohmeda Aespire Carestation. This device is not marketed in the US. | Datex-Ohmeda S/5 Aespire Anesthesia System 7100 | 2 | 10/11/2005 | 10/12/2005 | 04/01/2007 | 2183066 | Datex-Ohmeda Inc | Two conditions may occur with the S/5 Aespire Anesthesia System. The internal auxiliary flowmeter tubing may become disconnected if the auxiliary oxygen supply tubing is occluded and the patient may not receive desired supplemental oxygen which could lead to serious injury if undetected. Also the O2 switch may leak or fail.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=42271 | Z-0037-06 | Datex-Ohmeda S/5 Aespire Anesthesia System 7100 | Datex-Ohmeda S/5 Aespire Anesthesia System 7100 | 2 | 10/11/2005 | 10/12/2005 | 04/01/2007 | 2183066 | Datex-Ohmeda Inc | Two conditions may occur with the S/5 Aespire Anesthesia System. The internal auxiliary flowmeter tubing may become disconnected if the auxiliary oxygen supply tubing is occluded and the patient may not receive desired supplemental oxygen which could lead to serious injury if undetected. Also the O2 switch may leak or fail.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121287 | Z-0008-2014 | Draeger Fabius GS Premium, Fabius OS, Fabius Tiro, Fabius Tiro D-M Product Usage: used to administer anesthesia and ventilation to patients during surgical procedures. | Draeger Fabius | 1 | 10/10/2013 | 10/10/2013 | 09/23/2014 | 2517967 | Draeger Medical, Inc. | During final testing of the Draeger Fabius GS Premium, Fabius OS, and Fabius Tiro anesthesia machines, some units did not pass the high voltage test and the minimum clearance between an electrical component and the unit housing was not maintained for some from a specific batch.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=112797 | Z-0018-2013 | Ultra PC% Cabinet Mount Package G, Model #36600, Serial # 13593, 13594, 13595, 13596, 13597, 13598, 13599, 13600, 13602, 13605, 13610, 13611, 13612, 13615, 13616, 13617, 13621, 13622, 13624, 13625, 13627, 13628, 13633, 13634, 13635, 13636, 13637, 13638, 13639, 13640, 13642, 13643, 13644, 13647, 13650, 13669, 13675, 13673, 13678, 13679, 13680, 13683, 13685, 13687, 13691, 13693, 13694, 13695, 13696, 13697, 13701, 13702, 13705, 13706, 13709, 13710, 13711, 13724, 13725, 13728, 13729, 13735, 13739, 13743, 13746, 13750, 13751, 13753, 13754, 13755, 13756, 13760, 13762, 13766, 13774, 13775, 13777, 13781, 13782, 13783, 13785, 13786, 13788, 13792, 13794, 13796, 13799, 13800, 13803, 13804, 13809, 13810, 28298, 28299, 28304, 28305, 28306, 28307, Ultra PC% Cabinet Mount Pkg I, Model #36800, Serial # 13601, 13603, 13607, 13608, 13609, 13613, 13618, 13620, 13623, 13626, 13629, 13630, 13631, 13632, 13641, 13645, 13646, 13648, 13649, 13651, 13652, 13653, 13654, 13655, 13656, 13657, 13658, 13659, 13660, 13661, 13665, 13666, 13667, 13668, 13671, 13672, 13674, 13676, 13677, 13684, 13688, 13689, 13698, 13700, 13707, 13723, 13727, 13730, 13732, 13734, 13741, 13742, 13757, 13758, 13759, 13761, 13762, 13763, 13764, 13765, 13767, 13769, 13789, 13790, 13791, 13795, 13797, 13801, 13802, 13806, 13808, 28300, 28302, 28303, Ultra PC% Cabinet Mount Pkg G, Model#36600-NS, Serial# 13805, 13807, 13811, 13812. Ultra PC% Cabinet Mount Pkg H, Model#36800-NS, Serial# 13744, 13745,13752, 13773, 13774, 13778, 13779, 13780 Product Usage: Conscious Sedation Unit Analgesia Gas Machine. | Ultra PC% Cabinet Mount Package G, | 1 | 10/09/2012 | 10/09/2012 | 02/01/2013 | 2020813 | Accutron Inc | Accutron is recalling the Flowmeter (Analgesia Gas Machine) because when it flows Nitrous Gas without Oxygen gas being turned on, it functions in an incorrect manner.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=112796 | Z-0017-2013 | Flow Meter, CM Assembly, Model #27025, Serial # 13686, 13690, 13712, 13793. Flowmeter, CM Assembly, Model#31970, Serial#13692. Product Usage: Conscious Sedation Unit Analgesia Gas Machine. | Flow Meter, CM Assembly | 1 | 10/09/2012 | 10/09/2012 | 02/01/2013 | 2020813 | Accutron Inc | Accutron is recalling the Flowmeter (Analgesia Gas Machine) because when it flows Nitrous Gas without Oxygen gas being turned on, it functions in an incorrect manner.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=112799 | Z-0020-2013 | Newport Portable Package, Model #50000, Serial # 13591, 13592, 13604, 13606, 13614, 13662, 13663, 13664, 13726, 13731, 13733, 13736, 13737, 13738, 13740, 13768, 13770, 13771, 13772, 13784, 13787. Product Usage: Portable Analgesia Gas Machine. | Newport Portable Package | 1 | 10/09/2012 | 10/09/2012 | 02/01/2013 | 2020813 | Accutron Inc | Accutron is recalling the Flowmeter (Analgesia Gas Machine) because when it flows Nitrous Gas without Oxygen gas being turned on, it functions in an incorrect manner.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=112798 | Z-0019-2013 | CM Flowmeter Pkg H, Model #36700, Serial # 13749 Product Usage: Conscious Sedation Unit Analgesia Gas Machine. | CM Flowmeter Pkg H | 1 | 10/09/2012 | 10/09/2012 | 02/01/2013 | 2020813 | Accutron Inc | Accutron is recalling the Flowmeter (Analgesia Gas Machine) because when it flows Nitrous Gas without Oxygen gas being turned on, it functions in an incorrect manner.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=73333 | Z-0087-2009 | Draeger Medical Apollo Anesthesia Machine, catalog number 8606500 | Apollo Anesthesia Machine | 2 | 10/08/2008 | 10/08/2008 | 09/01/2009 | 2517967 | Draeger Medical, Inc. | Sporadic errors in various device functions, including low readings of the single gas flows for oxygen, nitrous oxide, or air delivery, or mechanical ventilation failure. Alarms and error messages function as designed.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94055 | Z-0009-2011 | GE Datex-Ohmeda Avance Anesthesia System, manufactured by GE Healthcare, P.O. Box 900, FIN-00031 GE, Finland | GE Datex-Ohmeda Avance Anesthesia System | 2 | 10/06/2010 | 10/06/2010 | 09/13/2012 | 2126677 | GE Healthcare, LLC | Users may inadvertently leave a test plug in the inspiratory flow sensor after completing Low P Leak Check on the SCGO (Switched Common Gas Outlet) configuration of Avance, Amingo, and Aisys anesthesia machines that may impact patient safety. If the test plug remains in place and therapy is initiated, the machine will not provide gas to the patient. This would inhibit mechanical or manual ventila
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94054 | Z-0008-2011 | GE Datex-Ohmeda Aisys Anesthesia System, manufactured by GE Healthcare, P.O. Box 900, FIN-00031 GE, Finland | GE Datex-Ohmeda Aisys Anesthesia System | 2 | 10/06/2010 | 10/06/2010 | 09/13/2012 | 2126677 | GE Healthcare, LLC | Users may inadvertently leave a test plug in the inspiratory flow sensor after completing Low P Leak Check on the SCGO (Switched Common Gas Outlet) configuration of Avance, Amingo, and Aisys anesthesia machines that may impact patient safety. If the test plug remains in place and therapy is initiated, the machine will not provide gas to the patient. This would inhibit mechanical or manual ventila
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94057 | Z-0010-2011 | GE Datex-Ohmeda Amingo Anesthesia System, manufactured by GE Healthcare, P.O. Box 900, FIN-00031 GE, Finland | GE Datex-Ohmeda Amingo Anesthesia System | 2 | 10/06/2010 | 10/06/2010 | 09/13/2012 | 2126677 | GE Healthcare, LLC | Users may inadvertently leave a test plug in the inspiratory flow sensor after completing Low P Leak Check on the SCGO (Switched Common Gas Outlet) configuration of Avance, Amingo, and Aisys anesthesia machines that may impact patient safety. If the test plug remains in place and therapy is initiated, the machine will not provide gas to the patient. This would inhibit mechanical or manual ventila
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=73230 | Z-0020-2009 | GE Datex-Ohmeda Aisys, Datex-Ohmeda, Inc., a General Electric Company going to market as GE Healthcare, GE Healthcare PO Box 7550, Madison, WI. 53707-7550 Datex-Ohmeda Aisys Intended to provide general inhalation anesthesia and ventilatory support to a wide range of patients. The device is intended for volume or pressure control ventilation. The Aisys is not suitable for use in a MRI environment. | GE Datex-Ohmeda Aisys | 2 | 10/06/2008 | 10/06/2008 | 03/08/2016 |  | Datex - Ohmeda, Inc | The electronic vaporization system in the GE Aisys Anesthesia machine contains the following components: Backpressure Valve, Inflow Check Valve and Cassette Interface Board, that can contribute to independent failures.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=215848 | Z-2617-2025 | Brand Name: Atlan Product Name: Atlan A350XL Model/Catalog Number: 8621600 Software Version: Not Applicable Product Description: The Atlan is an anesthesia workstation intended for use in anesthetizing adults, pediatrics, and neonates. The device can be used for mechanical ventilation, manual ventilation, pressure-supported spontaneous breathing, and spontaneous breathing. The device is specified for inhalational anesthesia and/or patient ventilation in accordance with the intended use during surgical or diagnostic interventions Component: Not applicable | Atlan | 1 | 10/01/2025 | 10/01/2025 |  | 2517967 | Draeger, Inc. | The device either indicated a failure of the piston ventilator before use or suffered a failure of the mechanical ventilation during use.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=215847 | Z-2616-2025 | Brand Name: Atlan Product Name: Atlan A350 Model/Catalog Number: 8211500 Software Version: Not applicable Product Description: The Atlan is an anesthesia workstation intended for use in anesthetizing adults, pediatrics, and neonates. The device can be used for mechanical ventilation, manual ventilation, pressure-supported spontaneous breathing, and spontaneous breathing. The device is specified for inhalational anesthesia and/or patient ventilation in accordance with the intended use during surgical or diagnostic interventions Component: Not applicable | Atlan | 1 | 10/01/2025 | 10/01/2025 |  | 2517967 | Draeger, Inc. | The device either indicated a failure of the piston ventilator before use or suffered a failure of the mechanical ventilation during use.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188580 | Z-2595-2021 | Flow sensors kept as stand-alone user replaceable spare parts for the GEHC anethesia machines: 2089610-001 FLOW SENSOR, LEGACY VAR ORF BCG (blue, cleanable) 2089610-001-S FLOW SENSOR, LEGACY VAR ORF BCG, SERVICE (blue, cleanable) 2087640-001 FLOW SENSOR, LEGACY VAR ORF AUTOCLAVABLE BCG (gray, autoclavable) 2087640-001-S FLOW SENSOR, LEGACY VAR ORF AUTOCLAVABLE BCG, SERVICE (gray, autoclavable) 2096513-001-S FLOW SENSOR ASSEMBLY 5697309 R-FMI34109-FLOW SENSOR, LEGACY VAR ORF BCG 5697310 R-FMI34109-FLOW SENSOR, LEGACY VAR ORF AUTOCLAVABLE BCG | Aestiva MRI | 2 | 09/30/2021 |  |  | 2126677 | GE Healthcare, LLC | The flow sensors built prior to June 2021 could have damaged tubes with small punctures or cuts that could cause leaks resulting in incorrect anesthesia machine tidal volumes potentially leading to over-delivery of volume to the patient.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=188592 | Z-2596-2021 | Flow sensors that may be installed in the following GEHC anesthesia machines: (1) Aestiva MRI, GTIN 0080682102339; (2) Aestiva 7900; (3) Aestiva 7100; (4) Aespire System Carestation, S/5 Protiva Config., Aespire 100 Compact Wall Mount (Aespire 7100); (5) S/5 Aespire 7900; (6) Aespire View; (7) Avance; (8) Avance CS2, GTIN 00840682102292; (9) Aisys; (10) Aisys CS2, GTIN 00840682102322; (11) Amingo; (12) 9100 NXT Config, Aelite NXT; (13) Protiva 7100; | Aestiva 7900 | 2 | 09/30/2021 |  |  | 2126677 | GE Healthcare, LLC | The flow sensors built prior to June 2021 could have damaged tubes with small punctures or cuts that could cause leaks resulting in incorrect anesthesia machine tidal volumes potentially leading to over-delivery of volume to the patient.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94224 | Z-2478-2010 | Fabius GS Anesthesia Machine, OR/Anesthesia, Draeger Medical, Inc., 3135 Quarry Road, Telford, PA 18969 inhalation anesthesia machine for use in operating, induction and recovery rooms | Fabius GS | 2 | 09/23/2010 | 09/23/2010 | 08/23/2011 | 2510954 | Draeger Medical Systems, Inc. | potential failures to Electri-Cord Manufacturing AC power cord
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94225 | Z-2479-2010 | Fabius Tiro Anesthesia Machine, OR/Anesthesia, Draeger Medical, Inc., 3135 Quarry Road, Telford, PA 18969 Compact anesthesia system that offers the full spectrum of anesthesia ventilation to military anesthesiologists/ CRNAs in the field of operations. Can be used in a variety of mobile and stationary military applications where general anesthesia is required. | Fabius Tiro | 2 | 09/23/2010 | 09/23/2010 | 08/23/2011 | 2510954 | Draeger Medical Systems, Inc. | potential failures to Electri-Cord Manufacturing AC power cord
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=72443 | Z-2338-2008 | Fabius Trio Gas machine, anesthesia or analgesia. Catalog number 8606000. The product is used in operating rooms and ER, for patient intubation. | Fabius Trio Gas machine | 2 | 09/23/2008 | 09/23/2008 | 03/08/2010 | 2517967 | Draeger Medical, Inc. | The parent company became aware of instances where the lower side rail of the frame of the machine broke inwards when being moved across a threshold; resulting in the machine being unstable.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=222077 | Z-3192-2026 | Et Control, Software option for anesthesia gas machine to achieve and maintain targeted end tidal oxygen and anesthetic agents, Field replaceable unit (FRU) part numbers: 1011-8000-000-S, 5700665 | Et Control | 2 | 09/16/2026 | 09/16/2026 |  | 2112667 | Datex-Ohmeda, Inc. | Anesthesia Delivery Systems containing certain electronic gas mixer assemblies can automatically transition to Alternate O2 control. In Alternate O2 control, the system delivers 100% O2 to the breathing system enabling continued controlled or manual ventilation and initiates audible and visual alarms. When this issue occurs, anesthetic agent delivery will be interrupted and requires clinician intervention. If this is not provided, the patient may experience decreased anesthetic agent effect.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=222074 | Z-3189-2026 | Aisys, Datex-Ohmeda anesthesia system, Part Number 1011-9050-000 | Aisys | 2 | 09/16/2026 | 09/16/2026 |  | 2112667 | Datex-Ohmeda, Inc. | Anesthesia Delivery Systems containing certain electronic gas mixer assemblies can automatically transition to Alternate O2 control. In Alternate O2 control, the system delivers 100% O2 to the breathing system enabling continued controlled or manual ventilation and initiates audible and visual alarms. When this issue occurs, anesthetic agent delivery will be interrupted and requires clinician intervention. If this is not provided, the patient may experience decreased anesthetic agent effect.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=222076 | Z-3191-2026 | Aisys CS2 with Et Control, Datex-Ohmeda anesthesia system, Part Number 1011-9000-000 | Aisys CS2 with Et Control | 2 | 09/16/2026 | 09/16/2026 |  | 2112667 | Datex-Ohmeda, Inc. | Anesthesia Delivery Systems containing certain electronic gas mixer assemblies can automatically transition to Alternate O2 control. In Alternate O2 control, the system delivers 100% O2 to the breathing system enabling continued controlled or manual ventilation and initiates audible and visual alarms. When this issue occurs, anesthetic agent delivery will be interrupted and requires clinician intervention. If this is not provided, the patient may experience decreased anesthetic agent effect.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=222075 | Z-3190-2026 | Aisys CS2, Datex-Ohmeda anesthesia system, Part Number 1011-9055-000 | Aisys CS2 | 2 | 09/16/2026 | 09/16/2026 |  | 2112667 | Datex-Ohmeda, Inc. | Anesthesia Delivery Systems containing certain electronic gas mixer assemblies can automatically transition to Alternate O2 control. In Alternate O2 control, the system delivers 100% O2 to the breathing system enabling continued controlled or manual ventilation and initiates audible and visual alarms. When this issue occurs, anesthetic agent delivery will be interrupted and requires clinician intervention. If this is not provided, the patient may experience decreased anesthetic agent effect.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34080 | Z-1441-04 | Anestar Anesthesia Delivery System | Anestar | 2 | 09/08/2004 | 09/09/2004 | 10/03/2006 | 2221819 | Datascope Corp | Failure of the Anestar AC Main Power Switch. Battery should last 30 minutes, then ventilation and monitoring will fail.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121105 | Z-2161-2013 | Tec 6 continuous Flow Vaporizer and Tec 6 Plus Anesthesia Vaporizer. The Datex-Ohmeda Tec 6 Plus and Tec 6 Plus NAD Variant is an electronic vaporizer which delivers the anesthetic agent desflurane. The Tec 6 Plus attaches only to Datex-Ohmeda anesthesia machines equipped with the proprietary Selectatec manifold. The Tec 6 Plus NAD Variant attaches only to NAD anesthesia machines equipped with the NAD exclusion type manifold. | Tec 6 Plus Vaporizer | 2 | 09/06/2013 | 09/06/2013 | 04/17/2015 | 2126677 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the pre-operative check, low pressure leak test at 1% dial setting of the Tec 6 and Tec 6 Plus Vaporizers. The user manual pre-operative check, low pressure leak test may not detect the full range of leaks from seal wear degradation in the vaporizers. If the leak is not detected, it will result in fresh gas leaking to
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=53647 | Z-1203-2007 | Kion Anesthesia Workstation Maquet Critical Care AB SE-171-95 Solna, Sweden | Kion Anesthesia Workstation | 2 | 09/06/2007 | 09/07/2007 | 12/07/2007 | 3008355164 | MAQUET Inc. | Overcharging of batteries and the release of acid fumes
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209456 | Z-2951-2024 | Isoflurane E_Vap.(Safety Filling adapter): used with the A9 Anesthesia System-Used to provide the vapor of an anesthetic agent with controllable concentration to an A9 anesthesia system. Part Number: Part Number 115-066758-00 | Mindray | 2 | 09/04/2024 | 09/04/2024 |  | 2221819 | Mindray DS USA, Inc. dba Mindray North America | V90 electronic vaporizer anesthetic agent output may be higher than the set amount due to a component within the vaporizer causing an incomplete seal in some cases. Should this issue occur, the A9 Anesthesia System will generate alarms associated with abnormal vaporizer output. If the anesthetic concentration is too high, the patient can inhale too much anesthetic agent; the patient may suffer from symptoms such as vertigo and drowsiness after the operation. In severe cases, the patient may suffer from neurological sequelae such as hypomnesia and inattention
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209457 | Z-2952-2024 | Desflurane E_Vap.(Saf-T-Fill) used with the A9 Anesthesia System-Used to provide the vapor of an anesthetic agent with controllable concentration to an A9 anesthesia system. Part Number: Part Number115-066759-00 | Mindray | 2 | 09/04/2024 | 09/04/2024 |  | 2221819 | Mindray DS USA, Inc. dba Mindray North America | V90 electronic vaporizer anesthetic agent output may be higher than the set amount due to a component within the vaporizer causing an incomplete seal in some cases. Should this issue occur, the A9 Anesthesia System will generate alarms associated with abnormal vaporizer output. If the anesthetic concentration is too high, the patient can inhale too much anesthetic agent; the patient may suffer from symptoms such as vertigo and drowsiness after the operation. In severe cases, the patient may suffer from neurological sequelae such as hypomnesia and inattention
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209455 | Z-2950-2024 | Sevoflurane E_Vap.(Safety Filling adapter) used with the A9 Anesthesia System-Used to provide the vapor of an anesthetic agent with controllable concentration to an A9 anesthesia system. Part Number: 115-066756-00 | Mindray | 2 | 09/04/2024 | 09/04/2024 |  | 2221819 | Mindray DS USA, Inc. dba Mindray North America | V90 electronic vaporizer anesthetic agent output may be higher than the set amount due to a component within the vaporizer causing an incomplete seal in some cases. Should this issue occur, the A9 Anesthesia System will generate alarms associated with abnormal vaporizer output. If the anesthetic concentration is too high, the patient can inhale too much anesthetic agent; the patient may suffer from symptoms such as vertigo and drowsiness after the operation. In severe cases, the patient may suffer from neurological sequelae such as hypomnesia and inattention
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209454 | Z-2949-2024 | Sevoflurane E_Vap.(Quik-Fil):used with the A9 Anesthesia System-Used to provide the vapor of an anesthetic agent with controllable concentration to an A9 anesthesia system. Part Number 115-066755-00 | Mindray | 2 | 09/04/2024 | 09/04/2024 |  | 2221819 | Mindray DS USA, Inc. dba Mindray North America | V90 electronic vaporizer anesthetic agent output may be higher than the set amount due to a component within the vaporizer causing an incomplete seal in some cases. Should this issue occur, the A9 Anesthesia System will generate alarms associated with abnormal vaporizer output. If the anesthetic concentration is too high, the patient can inhale too much anesthetic agent; the patient may suffer from symptoms such as vertigo and drowsiness after the operation. In severe cases, the patient may suffer from neurological sequelae such as hypomnesia and inattention
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=166166 | Z-2956-2018 | GE 9100c NXT anesthesia systems | GE 9100c NXT | 2 | 08/30/2018 |  | 12/16/2020 | 3003513129 | GE Medical Systems China Co., Ltd. | On systems equipped with the Auxiliary Common Gas Outlet (ACGO), there is a potential for the auxiliary gas switch to unexpectedly switch from the ACGO position back to the circle position. On systems equipped with the Aux O2 + Air function the switch could unexpectedly switch from the Aux O2 + Air position back to the circle position. These situations could lead to hypoxia.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=166164 | Z-2954-2018 | GE Carestation 620/650/650c, A1 and A2 (non-US distribution) anesthesia systems | GE Carestation 620/650/650c | 2 | 08/30/2018 |  | 12/16/2020 | 3003513129 | GE Medical Systems China Co., Ltd. | On systems equipped with the Auxiliary Common Gas Outlet (ACGO), there is a potential for the auxiliary gas switch to unexpectedly switch from the ACGO position back to the circle position. On systems equipped with the Aux O2 + Air function the switch could unexpectedly switch from the Aux O2 + Air position back to the circle position. These situations could lead to hypoxia.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=166165 | Z-2955-2018 | GE Carestation 30 anesthesia systems | GE Carestation 30 | 2 | 08/30/2018 |  | 12/16/2020 | 3003513129 | GE Medical Systems China Co., Ltd. | On systems equipped with the Auxiliary Common Gas Outlet (ACGO), there is a potential for the auxiliary gas switch to unexpectedly switch from the ACGO position back to the circle position. On systems equipped with the Aux O2 + Air function the switch could unexpectedly switch from the Aux O2 + Air position back to the circle position. These situations could lead to hypoxia.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148395 | Z-2689-2016 | A7 Anesthesia Delivery System, a device used to administer to a patient, continuously or intermittently, a general inhalation anesthetic and to maintain a patient's ventilation. Part numbers 0632F-PA0000X (US domestic) and 0632B-00014 (international) | Mindray | 2 | 08/30/2016 |  | 03/28/2017 | 2221819 | Mindray DS USA, Inc. dba Mindray North America | Potential for a leak to occur on the back-up O2 and air e-size cylinder yokes on the A7 Anesthesia Delivery System.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=148663 | Z-2692-2016 | UniCel DxH Slidemaker Stainer Coulter Cellular Analysis System, Catalog No. 775222 UniCel DxH Slidemaker Stainer (DxH SMS) Instructions for Use (IFU) Part Number B26647AC The DxH Slidemaker Stainer is a fully automated slide preparation and staining device that aspirates a whole-blood sample, smears a blood film on a clean microscope slide, and delivers a variety of fixatives, stains, buffers, and rinse solutions to that blood smear | UniCel DxH Slidemaker Stainer Coulter Cellular Analysis System | 2 | 08/30/2016 |  | 10/06/2016 | 2050012 | Beckman Coulter Inc. | Beckman Coulter is recalling the UniCel DxH Slidemaker Stainer Coulter Cellular Analysis System to reduce the potential for flammable liquids coming into contact with electrical components in the Stainer module possibly causing a fire.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=28387 | Z-0937-03 | Datex-Ohmeda S/5 Anesthesia Delivery Units | Datex-Ohmeda | 2 | 08/27/2003 | 08/28/2003 | 11/08/2005 | 2183066 | Datex-Ohmeda Inc | Possible disruption of mechanical ventilation or disruption of anesthetic agent delivery.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93109 | Z-2263-2010 | Philips M1013A IntelliVue G1 Anesthesia Gas Modules with manifolds | Philips M1013A IntelliVue G1 Anesthesia Gas Modules with manifolds | 2 | 08/24/2010 | 08/24/2010 | 03/09/2012 | 1218950 | Philips Healthcare Inc. | The manifold seal in the Philips M1657B Water Trap used on the Philips M1013A IntelliVue G1, M1019A, IntelliVue G5, and Anesthesia Gas Modules M1026B/BT may be unintentionally removed when changing the traps. A missing seal may result in leakage and may cause incorrect monitor readings of inspiratory or expiratory gas levels.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93111 | Z-2265-2010 | Philips M1026B IntelliVue Anesthesia Gas Modules with manifolds | Philips M1026B IntelliVue Anesthesia Gas Modules with manifolds | 2 | 08/24/2010 | 08/24/2010 | 03/09/2012 | 1218950 | Philips Healthcare Inc. | The manifold seal in the Philips M1657B Water Trap used on the Philips M1013A IntelliVue G1, M1019A, IntelliVue G5, and Anesthesia Gas Modules M1026B/BT may be unintentionally removed when changing the traps. A missing seal may result in leakage and may cause incorrect monitor readings of inspiratory or expiratory gas levels.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93112 | Z-2266-2010 | Philips M1026BT Gas Modules with manifolds | Philips M1026BT Gas Modules with manifolds | 2 | 08/24/2010 | 08/24/2010 | 03/09/2012 | 1218950 | Philips Healthcare Inc. | The manifold seal in the Philips M1657B Water Trap used on the Philips M1013A IntelliVue G1, M1019A, IntelliVue G5, and Anesthesia Gas Modules M1026B/BT may be unintentionally removed when changing the traps. A missing seal may result in leakage and may cause incorrect monitor readings of inspiratory or expiratory gas levels.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93110 | Z-2264-2010 | Philips M1019A IntelliVue G5 Anesthesia Gas Modules with manifolds | Philips M1019A IntelliVue G5 Anesthesia Gas Modules with manifolds | 2 | 08/24/2010 | 08/24/2010 | 03/09/2012 | 1218950 | Philips Healthcare Inc. | The manifold seal in the Philips M1657B Water Trap used on the Philips M1013A IntelliVue G1, M1019A, IntelliVue G5, and Anesthesia Gas Modules M1026B/BT may be unintentionally removed when changing the traps. A missing seal may result in leakage and may cause incorrect monitor readings of inspiratory or expiratory gas levels.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=221405 | Z-2860-2026 | Flow Anesthesia Machine Flow-i C20 6677200 Flow-i C30 6677300 Flow-i C40 6677400 Flow-i C20 6888520 Flow-i C30 6888530 Flow-i C40 6888540 Flow-c 6887700 Flow-e 6887900 | Flow Anesthesia Machine | 2 | 08/03/2026 | 08/03/2026 |  | 3002808128 | Maquet Critical Care AB | Events observed where anesthetic agent was not delivered when the vaporizer was activated. Failure to deliver the anesthetic agent may result in the patient receiving insufficient anesthesia and may lead to intraoperative awareness.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=77328 | Z-1635-2009 | Datex-Ohmeda S/5 Anesthesia Delivery Unit Intended to provide general inhalation anaesthesia and ventilatory support to a wide range of patients. Intended for volume or pressure control, pressure support and synchronized intermittent mandatory (SIMV) ventilation modes. | Datex-Ohmeda S/5 Anesthesia Delivery Unit | 2 | 07/30/2009 | 07/30/2009 | 05/17/2012 |  | GE Healthcare | GE Healthcare has received reports of breathing circuit or tubing misconnections of the Anesthesia Delivery Unit (ADU), that may impact patient safety.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=77327 | Z-1634-2009 | Datex-Ohmeda AS/3 Anesthesia Delivery Unit An Anesthesia Gas Machine for composition of anesthetic breathing gas and manual or mechanical ventilation for patients undergoing anesthesia. Indicated for all persons that require inhalation anesthesia. | Datex-Ohmeda AS/3 Anesthesia Delivery Unit | 2 | 07/30/2009 | 07/30/2009 | 05/17/2012 |  | GE Healthcare | GE Healthcare has received reports of breathing circuit or tubing misconnections of the Anesthesia Delivery Unit (ADU), that may impact patient safety.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=77325 | Z-1633-2009 | GE Datex-Ohmeda ADU Carestation Intended to provide general inhalation anesthesia and ventilatory support to a wide range of patients. | GE Datex-Ohmeda ADU Carestation | 2 | 07/30/2009 | 07/30/2009 | 05/17/2012 |  | GE Healthcare | GE Healthcare has received reports of breathing circuit or tubing misconnections of the Anesthesia Delivery Unit (ADU), that may impact patient safety.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92556 | Z-2113-2010 | BleaseSirius Anaesthesia System, gas machine. Blease Beech House, Chiltern Court Asheridge Road, Chesham, Bucks HP5 2PX, England Intended for use in the hospital environment and operating room. | BleaseSirius Anaesthesia Machine | 2 | 07/29/2010 | 07/29/2010 | 12/03/2010 | 3010157426 | Spacelabs Healthcare, Llc | Potential for the touch screen on the ventilator portion of the device to go blank when touched. Possibility that this failure may also cause mechanical ventilation to stop.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138314 | Z-2271-2015 | Fabius MRI Anesthesia Machine; Fabius MRI is an inhalation anesthesia machine for use in MRI environments in operating, induction and recovery rooms. | Fabius MRI Anesthesia Machine | 2 | 07/27/2015 | 07/27/2015 | 05/05/2017 | 2517967 | Draeger Medical, Inc. | the Fabius MRI or parts of the system were attracted by the magnetic field of the MRI. If the Fabius MRI is positioned too close to the MRI, its parts can be loosened as a result of the magnetic force of the MRI.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=194360 | Z-1427-2022 | GE Healthcare Avance CS2, Avance CS2 Pro Anesthesia Systems, Part #1009-9050-000; and Aisys CS2 Anesthesia Systems, Part # 1011-9050-000. | GE Healthcare Avance and Aisys | 2 | 07/20/2022 |  |  | 2126677 | GE Healthcare, LLC | The base of the anesthesia system can have a crack and could potentially fracture resulting in a tip or overbalance of the anesthesia system. This can result in potential injury if it falls on a person.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=194361 | Z-1389-2022 | GE Healthcare Aisys CS2 Anesthesia System, Part number 1011-9050-000; and Avance CS2 and Avance CS2 Pro Anesthesia System, Part number 1009-9050-000. | GE Healthcare Aisys and Avance | 1 | 07/19/2022 | 07/19/2022 | 10/21/2024 | 2126677 | GE Healthcare, LLC | There is a potential reversal of the O2 and air cylinder pressure transducer connections.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=134486 | Z-2083-2015 | The FLOW-i Anesthesia System Axillary O2 and Suction Module, Device Part number: 66 79 847 FLOW-i Anesthesia System: Is indicated for administering inhalation anesthesia while controlling the entire ventilation of patients with no ability to breathe, as well as in supporting patients with a limited ability to breathe.The system is intended for use in a wide range of population: from neonatal to adult patients. The system is intended for use in hospital environments, except for the MRI environment, by healthcare professionals trained in inhalation anesthesia administration. Auxiliary O2 and Suction Module: Is an option available for the FLOW-i Anesthesia System. The module contains two main parts: Flowmeter unit: The intended use of the flowmeter is to provide oxygen for patient therapy. Suction unit: The intended use is to extract body fluids from the stomach and airways. | FLOW-i | 2 | 07/15/2015 | 07/15/2015 | 04/29/2016 | 3008355164 | Maquet Cardiovascular Us Sales, Llc | Maquet has identified that in some FLOW-i Anesthesia System Auxiliary O2 and Suction Modules, the suctioning function cannot be activated or controlled by the suction unit on/off switch leading to a decrease or loss of suctioning capability of the modules.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=68298 | Z-1315-2008 | Fabius GS Anesthesia Machine; Gas machine for anesthesia or analgesia; Catalog Number: 8604700; Drager Medical Inc, Telford, PA 18969. | Fabius GS Anesthesia Machine | 2 | 07/15/2008 | 07/16/2008 | 09/09/2008 | 2517967 | Draeger Medical, Inc. | Detachment of Component: The chassis on some machines were manufactured out of specifications. As a result, the caster may break loose from the chassis.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=44224 | Z-1196-06 | Fresh Gas Control (A-FGC1) Unit, a component of the S/5 Anesthesia Delivery Unit. The Anesthesia Delivery Unit is designed to mix and dose respiratory gases and to ventilate the patient. The delivery unit is the basis for a flexible anesthesia system with full monitoring and data management capabilities. | S/5 Anesthesia Delivery Unit | 2 | 07/11/2006 | 07/12/2006 | 05/29/2007 | 2183066 | GE Healthcare | There exists the possibility that the ADU Fresh Gas Control unit (A-FGC1) can shut off during normal handling of the anesthesia machine and result in no anesthetic agent output.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=155543 | Z-2703-2017 | Carestation 620 A1, Carestation 650 A1, and Carestation 650c A1 | Carestation 620 A1, Carestation 650 A1, and Carestation 650c A1 | 2 | 07/07/2017 | 05/24/2017 | 02/13/2018 | 2126677 | GE Healthcare, LLC | Carestation 620, 650 and 650c has Potential for Elevated FiCO2 and Unexpected System Malfunction
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=109875 | Z-1925-2012 | FLOW-I Anesthesia System The system is intended for use in administrating anesthesia while controlling the entire ventilation for patients with no ability to breathe, as well as in supporting patients with a limited ability to breathe. | FLOW-I Anesthesia System | 1 | 07/06/2012 | 07/06/2012 | 09/09/2015 | 3008355164 | Maquet Cardiovascular Us Sales, Llc | A technical alarm may be generated on the FLOW-i system when using the MAN/AUTO switch to change ventilation modes from manual to automatic or from automatic to manual if the switch is not fully engaged in the "on" or "off" position but remains in the "in-between position.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193697 | Z-1286-2022 | Field replacement back up batteries distributed on or after April 1, 2019 for Avance CS2, Avance CS2 Pro, Avance, Amingo, Aespire View anesthesia devices | Avance CS2, Avance CS2 Pro, Avance, Amingo, Aespire View | 2 | 06/24/2022 |  |  | 2126677 | GE Healthcare, LLC | Backup batteries can fail earlier than their estimated life.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193549 | Z-1285-2022 | Back up batteries in Avance CS2 and Avance CS2 Pro anesthesia devices manufactured on or after April 1, 2019 | Avance | 2 | 06/24/2022 |  |  | 2126677 | GE Healthcare, LLC | Backup batteries can fail earlier than their estimated life.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193880 | Z-1279-2022 | Flow-e Anesthesia System : Intended for use in administering anesthesia while controlling the entire ventilation of patient Product Code/REF Number: 6887900 | Flow-e Anesthesia System | 1 | 06/23/2022 | 06/23/2022 | 08/06/2024 | 3013876692 | Getinge Usa Sales Inc | Potential for crack to form on the on/off switch of the suction unit, which may result in an unreasonable risk of harm to the patient. In the worst case, the switch can break, preventing the suction unit from being activated.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193879 | Z-1278-2022 | Flow-c Anesthesia System : intended for use in administering anesthesia while controlling the entire ventilation of patient Product Code/REF Number: 6887700 | Flow-c Anesthesia System | 1 | 06/23/2022 | 06/23/2022 | 08/06/2024 | 3013876692 | Getinge Usa Sales Inc | Potential for crack to form on the on/off switch of the suction unit, which may result in an unreasonable risk of harm to the patient. In the worst case, the switch can break, preventing the suction unit from being activated.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=79156 | Z-1425-2009 | GE Datex-Ohmeda Aisys, Anesthesia Gas-Machine. | Aisys | 2 | 06/23/2009 | 06/23/2009 | 12/10/2011 |  | GE Healthcare | Unintended shut-down: The anesthesia machine may start shutting down without human intervention due to a faulty On/Standby Switch. Prior to shutdown, the system will issue an alarm and play a message that shutdown will occur in 8 seconds. The failure only affects the electrical portion of the switch. If the clinician does not intervene, the patient could experience hypoventilation leading to h
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=80386 | Z-1426-2009 | GE Datex-Ohmeda Advance Anesthesia Gas-Machine. | Advance | 2 | 06/23/2009 | 06/23/2009 | 12/10/2011 |  | GE Healthcare | Unintended shut-down: The anesthesia machine may start shutting down without human intervention due to a faulty On/Standby Switch. Prior to shutdown, the system will issue an alarm and play a message that shutdown will occur in 8 seconds. The failure only affects the electrical portion of the switch. If the clinician does not intervene, the patient could experience hypoventilation leading to h
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164870 | Z-2252-2018 | Fabius GS MRI Anesthesia Machine; Cat. no. 8607300 Product Usage: Inhalation anesthesia machines for use in operating, induction and recovery rooms. They may be used with O2, N2O, and AIR supplied by a medical gas pipeline system or by externally mounted gas cylinders. Fabius series machines are equipped with a compact breathing system; providing fresh gas decoupling, PEEP, and pressure limitation. The following ventilation options are available: Volume Controlled Ventilation Pressure Controlled Ventilation Pressure Support (Optional) SIMV/PS (Optional) Manual Ventilation Spontaneous Breathing Fabius series anesthesia machines are equipped with an electrically driven and electronically controlled ventilator and monitors for airway pressure (P), volume (V), and inspiratory oxygen concentration (FiO2). As per IEC 60601-2-13 (Anesthetic Workstations and their Modules-Particular Requirements), additional monitoring of the concentrations of CO2 and anesthetic agent is required when the machine is in use. IEC 60601-2-13:2003 requires that a manual ventilation bag must be available for emergency use. Fresh gas enrichment is provided by the Drger Vapor anesthetic vaporizer. Fabius MRI can also be used in MRI scanner rooms with magnets 1.5 tesla and 3 tesla at a fringe field strength of 40 mtesla or less | Fabius GS MRI | 2 | 06/19/2018 |  | 05/12/2020 | 2517967 | Draeger Medical, Inc. | The Draeger anesthesia device may be able to dose 100% N2O. In the event of a fault, the S-ORC module would not prevent setting an N2O flow that would result in a hypoxic mixture from being dosed to the patient. Potential adverse outcomes include death of the patient.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164872 | Z-2254-2018 | Perseus A500 Anesthesia Machine; Cat. no. MK06000 Product V Product Usage: Intended for use in anesthetizing adults, children, and neonates and can be used for automatic and manual ventilation, pressuresupported spontaneous breathing, and spontaneous breathing. Perseus is equipped with airway monitoring, gas measurement and device monitoring, O2 insufflation, and an anesthetic gas receiving system. Anesthesia is achieved through a mixture of pure oxygen and Air (medical compressed air) or pure oxygen and nitrous oxide, with the addition of volatile anesthetic agents. Ventilation is accomplished on the patient through a laryngeal mask airway, a full-face mask, or an endotracheal tube. The integrated breathing system can be used with partial rebreathing (low-flow or minimum-flow). A non-rebreathing system such as the Kuhn or Medec Water System may be used at the external fresh-gas outlet (optional). Perseus is specified for inhalational anesthesia and/or patient ventilation in accordance with the intended use during surgical or diagnostic interventions. | Perseus A500 | 2 | 06/19/2018 |  | 05/12/2020 | 2517967 | Draeger Medical, Inc. | The Draeger anesthesia device may be able to dose 100% N2O. In the event of a fault, the S-ORC module would not prevent setting an N2O flow that would result in a hypoxic mixture from being dosed to the patient. Potential adverse outcomes include death of the patient.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164868 | Z-2250-2018 | Fabius GS Premium Anesthesia Machine; Cat. no. 8607000 Product Usage: Inhalation anesthesia machines for use in operating, induction and recovery rooms. They may be used with O2, N2O, and AIR supplied by a medical gas pipeline system or by externally mounted gas cylinders. Fabius series machines are equipped with a compact breathing system; providing fresh gas decoupling, PEEP, and pressure limitation. The following ventilation options are available: Volume Controlled Ventilation Pressure Controlled Ventilation Pressure Support (Optional) SIMV/PS (Optional) Manual Ventilation Spontaneous Breathing Fabius series anesthesia machines are equipped with an electrically driven and electronically controlled ventilator and monitors for airway pressure (P), volume (V), and inspiratory oxygen concentration (FiO2). As per IEC 60601-2-13 (Anesthetic Workstations and their Modules-Particular Requirements), additional monitoring of the concentrations of CO2 and anesthetic agent is required when the machine is in use. IEC 60601-2-13:2003 requires that a manual ventilation bag must be available for emergency use. Fresh gas enrichment is provided by the Drger Vapor anesthetic vaporizer. Fabius MRI can also be used in MRI scanner rooms with magnets 1.5 tesla and 3 tesla at a fringe field strength of 40 mtesla or less | Fabius GS Premium | 2 | 06/19/2018 |  | 05/12/2020 | 2517967 | Draeger Medical, Inc. | The Draeger anesthesia device may be able to dose 100% N2O. In the event of a fault, the S-ORC module would not prevent setting an N2O flow that would result in a hypoxic mixture from being dosed to the patient. Potential adverse outcomes include death of the patient.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164871 | Z-2253-2018 | Apollo Anesthesia Machine; Cat. no. 8605310 Product Usage: Indicated as a continuous flow anesthesia system. The Apollo may be used for manually assisted or automatic ventilation, delivery of gases and anesthetic vapor, and monitoring of oxygen and CO2 concentration, breathing pressure, respiratory volume, and anesthetic agent concentration and identification. It is an inhalation anesthesia machine for use in operating, induction, and recovery rooms. It can be used with rebreathing systems, semiclosed to virtually closed systems with low flow and minimal flow techniques, and non-rebreathing systems (with the Auxiliary Common Gas Outlet). It may be used with O2, N2O, and Air supplied by a medical gas pipeline system or by externally mounted gas cylinders. Anesthetic agent can be delivered via vaporizers mounted on the machine. The Apollo is equipped with a compact breathing system, providing fresh-gas decoupling, PEEP, and pressure limitation. It has an electrically driven and electronically controlled ventilator. | Apollo | 2 | 06/19/2018 |  | 05/12/2020 | 2517967 | Draeger Medical, Inc. | The Draeger anesthesia device may be able to dose 100% N2O. In the event of a fault, the S-ORC module would not prevent setting an N2O flow that would result in a hypoxic mixture from being dosed to the patient. Potential adverse outcomes include death of the patient.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=164869 | Z-2251-2018 | Fabius GS Tiro Anesthesia Machine; Cat. no. 8606000 Product Usage: Inhalation anesthesia machines for use in operating, induction and recovery rooms. They may be used with O2, N2O, and AIR supplied by a medical gas pipeline system or by externally mounted gas cylinders. Fabius series machines are equipped with a compact breathing system; providing fresh gas decoupling, PEEP, and pressure limitation. The following ventilation options are available: Volume Controlled Ventilation Pressure Controlled Ventilation Pressure Support (Optional) SIMV/PS (Optional) Manual Ventilation Spontaneous Breathing Fabius series anesthesia machines are equipped with an electrically driven and electronically controlled ventilator and monitors for airway pressure (P), volume (V), and inspiratory oxygen concentration (FiO2). As per IEC 60601-2-13 (Anesthetic Workstations and their Modules-Particular Requirements), additional monitoring of the concentrations of CO2 and anesthetic agent is required when the machine is in use. IEC 60601-2-13:2003 requires that a manual ventilation bag must be available for emergency use. Fresh gas enrichment is provided by the Drger Vapor anesthetic vaporizer. Fabius MRI can also be used in MRI scanner rooms with magnets 1.5 tesla and 3 tesla at a fringe field strength of 40 mtesla or less | Fabius GS Tiro | 2 | 06/19/2018 |  | 05/12/2020 | 2517967 | Draeger Medical, Inc. | The Draeger anesthesia device may be able to dose 100% N2O. In the event of a fault, the S-ORC module would not prevent setting an N2O flow that would result in a hypoxic mixture from being dosed to the patient. Potential adverse outcomes include death of the patient.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135941 | Z-1706-2015 | FLOW-i Anesthesia System The FLOW-i Anesthesia System is intended for use administering inhalation anesthesia on neonatal to adult patient populations. | Flow-i Anesthesia System | 1 | 06/08/2015 | 06/08/2015 | 07/06/2016 | 3008355164 | Maquet Cardiovascular Us Sales, Llc | In some Flow-i Anesthesia Systems, patient cassette can get dislodged which could result in gas leakage within the system internal circuit potentially leading to a stop in ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=220173 | Z-2207-2026 | Atlan A350. Model Number: 8211500. anesthesia workstation | Atlan anesthesia workstation | 1 | 06/04/2026 | 06/04/2026 |  | 2517967 | Draeger, Inc. | The Atlan anesthesia workstation either indicated a failure of the piston ventilator before use or suffered a failure of the mechanical ventilation during use.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=220174 | Z-2208-2026 | Atlan A350XL. Model Number: 8621600. anesthesia workstation | Atlan A350XL anesthesia workstation | 1 | 06/04/2026 | 06/04/2026 |  | 2517967 | Draeger, Inc. | The Atlan anesthesia workstation either indicated a failure of the piston ventilator before use or suffered a failure of the mechanical ventilation during use.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=118055 | Z-1443-2013 | Canister: CO2 canister with GE part numbers 1407-3200-000, 1407-7004-000, or M1084850. Intended to provide general inhalation anesthesia and ventilatory support. | Aespire, Aespire View, Avance, Avance CS2, Aisys, ADU, 9100, M900, M904, 9300 systems | 2 | 06/02/2013 | 06/03/2013 | 03/12/2014 | 2126677 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the reusable CO2 absorbent canister accessory used with Aespire, Aespire View, Avance, Avance CS2, Aisys, ADU, 9100, M900, M904, 9300 systems. Reusable CO2 absorbent canisters accessories may not seal properly.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=45203 | Z-0929-06 | Anestar and Anestar Plus Anesthesia System. | Anestar | 2 | 06/01/2006 | 06/02/2006 | 08/01/2006 | 2221819 | Datascope Corp | The Anestar and Anestar Plus Anesthesia System may produce periods of high Positive End-Expiratory Pressure (PEEP), in both manual and mechanical ventilation mode due to an issue with the Applied Pressure Limit (APL) valve.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=207330 | Z-1877-2024 | Vaporizer Sevoflurane Maquet Filling, Product code/REF number (6682282), Component of: Flow-i C20, Flow-i C30,Flow-i C40, Flow-c, Flow-e. | Vaporizer Sevoflurane Maquet Filling | 1 | 05/31/2024 | 05/31/2024 |  | 3013876692 | Getinge Usa Sales Inc | Sevoflurane used with the vaporizer may degrade to hydrogen fluoride, which could present a risk to patients or health care providers through inhalation and/or skin exposure. Potential harms include irritation of respiratory tract, which in worst case may lead to lung edema and/or severe hypocalcemia, blistering, superficial ulceration, and/or hypomagnesemia.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=154830 | Z-2093-2017 | Spacelabs Healthcare Arkon Anesthesia Delivery System is intended for use in the hospital environment and operating room. Absorber Adapter Assembly Service Kit, part number: 050-9045-00. Absorber Adapter Assembly, part number: 650-1648-01. Arkon Absorber Canister Carrier, part number: 380-0916-01. | Arkon Anesthesia Delivery System | 2 | 05/16/2017 |  | 10/16/2017 | 3010157426 | Del Mar Reynolds Medical, Ltd. | Reports of the Canister Carrier breaking during absorbent exchange or during anesthesia machine transit, causing a leak in the breathing circuit, which could cause hypoventilation in the patient.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90654 | Z-1527-2010 | GE, 1009-9002-000 Avance Anesthesia Machine and Monitor | GE Avance | 1 | 05/11/2010 | 05/11/2010 | 01/30/2012 | 2126677 | Ge Healthcare, Llc | Unanticipated shut-down: A specific lot of GE Healthcare Aisys and Avance machines control board wiring harness have a defect, which can cause the machine to unexpectedly shut down, terminating ventilation, anesthetic delivery, and potentially patient monitoring.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90650 | Z-1526-2010 | GE, 1011-9000-000 Aisys Anesthesia Machine and Monitor | GE Aisys | 1 | 05/11/2010 | 05/11/2010 | 01/30/2012 | 2126677 | Ge Healthcare, Llc | Unanticipated shut-down: A specific lot of GE Healthcare Aisys and Avance machines control board wiring harness have a defect, which can cause the machine to unexpectedly shut down, terminating ventilation, anesthetic delivery, and potentially patient monitoring.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=43016 | Z-0857-06 | Datex-Ohmeda S/5 Anesthesia Delivery Unit (ADU) | S/5 Aneshesia Delivery Unit | 2 | 05/08/2006 | 05/09/2006 | 01/23/2007 | 2183066 | GE Healthcare | Alarm sound-It is possible that the audible alarms from the speaker from the Datex-Ohmeda S/5 Anesthesia Delivery Unit (ADU) machine may dampen or disappear as a result of normal handling of the anesthesia machine. Visual alarms on the ADU are unaffected by this condition. The dampening or absence of audible alarms on the ADU may result in a delay in treatment.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=163355 | Z-1483-2018 | Fabius Tiro, anesthesia machine, catalog no. 8606000 Product Usage: The products are inhalation anesthesia machines for use in operating, induction and recovery rooms. They may be used with O2, N2O, and AIR supplied by a medical gas pipeline system or by externally mounted gas cylinders. Fabius series anesthesia machines are equipped with an electrically driven and electronically controlled ventilator and monitors for airway pressure (P), volume (V), and inspiratory oxygen concentration (FiO2). | Fabius Tiro | 1 | 05/07/2018 | 05/08/2018 | 04/16/2019 | 2517967 | Draeger Medical, Inc. | The automatic ventilation may fail if the position detection of the motor is disturbed. If the ventilator cover is damaged, manual ventilation may also fail.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=163356 | Z-1484-2018 | Fabius Tiro M, anesthesia machine, catalog no. 8608400 Product Usage: The products are inhalation anesthesia machines for use in operating, induction and recovery rooms. They may be used with O2, N2O, and AIR supplied by a medical gas pipeline system or by externally mounted gas cylinders. Fabius series anesthesia machines are equipped with an electrically driven and electronically controlled ventilator and monitors for airway pressure (P), volume (V), and inspiratory oxygen concentration (FiO2). | Fabius Tiro M | 1 | 05/07/2018 | 05/08/2018 | 04/16/2019 | 2517967 | Draeger Medical, Inc. | The automatic ventilation may fail if the position detection of the motor is disturbed. If the ventilator cover is damaged, manual ventilation may also fail.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=163353 | Z-1481-2018 | Fabius MRI, anesthesia machine, catalog no. 8607300 Product Usage: The products are inhalation anesthesia machines for use in operating, induction and recovery rooms. They may be used with O2, N2O, and AIR supplied by a medical gas pipeline system or by externally mounted gas cylinders. Fabius series anesthesia machines are equipped with an electrically driven and electronically controlled ventilator and monitors for airway pressure (P), volume (V), and inspiratory oxygen concentration (FiO2). This device can also be used in MRI scanner rooms with magnets 1.5 tesla and 3 tesla at a fringe field strength of 40 mtesla or less. | Fabius MRI | 1 | 05/07/2018 | 05/08/2018 | 04/16/2019 | 2517967 | Draeger Medical, Inc. | The automatic ventilation may fail if the position detection of the motor is disturbed. If the ventilator cover is damaged, manual ventilation may also fail.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=163354 | Z-1482-2018 | Fabius GS Premium, anesthesia machine, catalog no. 8607000 Product Usage: The products are inhalation anesthesia machines for use in operating, induction and recovery rooms. They may be used with O2, N2O, and AIR supplied by a medical gas pipeline system or by externally mounted gas cylinders. Fabius series anesthesia machines are equipped with an electrically driven and electronically controlled ventilator and monitors for airway pressure (P), volume (V), and inspiratory oxygen concentration (FiO2). | Fabius GS Premium | 1 | 05/07/2018 | 05/08/2018 | 04/16/2019 | 2517967 | Draeger Medical, Inc. | The automatic ventilation may fail if the position detection of the motor is disturbed. If the ventilator cover is damaged, manual ventilation may also fail.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213309 | Z-1631-2025 | GE Healthcare Carestation 650 A1, Model/REF Number 1012-9650-200. Anesthesia system. | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213305 | Z-1628-2025 | GE Healthcare CARESTATION 650 A1, Model/REF Number 1012-9650-000. Anesthesia system. | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213312 | Z-1634-2025 | GE Healthcare Carestation 650c A1, Model/REF Number 1012-9655-200. Anesthesia system. | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213311 | Z-1633-2025 | GE Healthcare CARESTATION 650C A2, Model/REF Number 1012-9655-002. Anesthesia system. | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213301 | Z-1624-2025 | GE Healthcare CARESTATION 620 A1, Model/REF Number 1012-9620-000. anesthesia systems | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213310 | Z-1632-2025 | GE Healthcare CARESTATION 650C A1, Model/REF Number 1012-9655-000. Anesthesia system. | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213313 | Z-1635-2025 | GE Healthcare CARESTATION 750 A1, Model/REF Number 1012-9750-000. Anesthesia system. | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213314 | Z-1636-2025 | GE Healthcare CARESTATION 750 A2, Model/REF Number 1012-9750-002. Anesthesia system. | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213302 | Z-1625-2025 | GE Healthcare CARESTATION 620 A2, Model/REF Number 1012-9620-002. Anesthesia system. | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213303 | Z-1626-2025 | GE Healthcare Carestation 620 SE, Model/REF Number 1012-9620-012. Anesthesia system. | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213315 | Z-1637-2025 | GE Healthcare CARESTATION 750c A1, Model/REF Number 1012-9755-000. Anesthesia system. | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213307 | Z-1630-2025 | GE Healthcare Carestation 650 SE, Model/REF Number 1012-9650-012. Anesthesia system. | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213316 | Z-1638-2025 | GE Healthcare CARESTATION 750c A2, Model/REF Number 1012-9755-002. Anesthesia system. | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213304 | Z-1627-2025 | GE Healthcare Carestation 620 A1, Model/REF Number 1012-9620-200. Anesthesia system. | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213306 | Z-1629-2025 | GE Healthcare CARESTATION 650 A2, Model/REF Number 1012-9650-002. Anesthesia system. | GE Healthcare CARESTATION | 1 | 04/29/2025 | 04/29/2025 |  | 3003513129 | GE Medical Systems China Co., Ltd. | Certain Carestation 620/650/650c and 750/750c Anesthesia Delivery Systems will not provide effective ventilation in Volume Control Ventilation (VCV) mode. In these systems, effective ventilation can be achieved in Pressure Control Ventilation (PCV) or Pressure Control Ventilation Volume Guarantee (PCV-VG) modes or with Manual ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144608 | Z-1562-2016 | GE Healthcare, Aisys and Aisys CS2. Intended to provide general inhalation anesthesia and ventilator support to a wide range of patients (neonatal, pediatric, and adult). | GE Datex-Ohmeda Aisys CS2 Anesthesia System | 2 | 04/22/2016 |  | 11/07/2016 | 2126677 | GE Medical Systems, LLC | The vacuum suction tubing could be kinked within the machine resulting in suction less than the 20 lpm required by ISO 10079-3. This could result in inadequate suction performance to clear the airway and inadequate ventilation of the patient.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=206777 | Z-1512-2024 | Drger Perseus A500 Anesthesia Workstation | Draeger Perseus A500 | 1 | 04/18/2024 | 04/18/2024 |  | 2517967 | Draeger Medical, Inc. | Unexpected shutdown while operating on battery power.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90223 | Z-1408-2010 | The BleaseSirius and BleaseFocus Anaesthesia Systems. | Anesthesia system | 2 | 04/16/2010 | 04/16/2010 | 08/23/2011 | 3002806541 | Blease Medical Equipment, Ltd. | Potential for delivery of a lower than set concentration of anaesthetic to a patient due to a sticking valve. The lower shuttle may become jammed, not permitting flow of anaesthesia gas through the vapouriser. The vapouriser back bar mounting valves will require exchange only on the systems identified in this matter, no other systems are affected.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=143994 | Z-1381-2016 | GE Healthcare, Aisys CS2 anesthesia.. | GE Datex-Ohmeda Aisys CS2 Anesthesia System | 2 | 04/13/2016 |  | 09/04/2018 | 2126677 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a potential safety issue involving a momentary, self-correcting anesthetic agent bolus when using 21% oxygen on all Aisys CS2 and upgraded Aisys anesthesia devices.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144046 | Z-1382-2016 | GE Healthcare, Aisys Carestation. | GE Healthcare, Aisys Carestation 2 | 2 | 04/13/2016 |  | 09/04/2018 | 2126677 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a potential safety issue involving a momentary, self-correcting anesthetic agent bolus when using 21% oxygen on all Aisys CS2 and upgraded Aisys anesthesia devices.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=89325 | Z-1313-2010 | Datascope AS3000 Anesthesia System; Mindray, North America. The AS3000 Anesthesia Delivery System is a continuous flow inhalation gas anesthesia system that delivers anesthetic vapor, provides for automatic and manual modes of ventilation, and is equipped with a monitoring system for ventilation, inspired and expired gas. The AS3000 is intended for use in operating rooms. It is used with O2, N2O and AIR supplied by a medical gas pipeline system or by externally mounted gas cylinders. Anesthetic agent can be delivered via vaporizers mounted on the machine. | NAVIGATOR ANESTHESIA DELIVERY SYSTEM, MODEL 0998-00-3024-01 | 2 | 04/08/2010 | 04/08/2010 | 04/25/2012 | 2221819 | Mindray DS USA, Inc., dba Datascope Patient Monitoring | An issue affecting the caster mount on the AS3000 System have been identified, specifically reports of the caster mount breaking, causing the wheel to fall off and the unit to tip.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=134476 | Z-1376-2015 | GE Healthcare, Aespire 7900, Aespire View Anesthesia Machines. Model numbers 1009-9012-000. Intended to provide general inhalation anesthesia and ventilatory support to a wide range of patients | Aespire 7900, Aespire View | 2 | 04/02/2015 | 04/02/2015 | 04/16/2015 | 2126677 | GE Healthcare | Specific Aespire 7900 Anesthesia Machines have an increased likelihood of failure of the CPU circuit board. This could result in a blank display or loss of mechanical ventilation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=162482 | Z-1199-2018 | DATEX-OHMEDA AISYS CS2 ALADIN2 CASSETTE Cassettes are intended to be used to deliver anesthetic agent. | Aisys CS2 Aladin Cassette | 2 | 03/26/2018 |  | 04/10/2020 | 2126677 | GE Healthcare, LLC | Non-Conforming devices were shipped from the service center without proper evaluation and therefore may be nonconforming.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=106127 | Z-1318-2012 | Anesthesia Machines (digital display assembly); A gas machine for anesthesia is a device used to administer to a patient, continuously or intermittently, a general inhalation anesthetic and to maintain a patient's ventilation. | Penlon Prima Anaesthesia Machines (SP2 and SP3) with Electronic Flowmeter Display | 2 | 03/26/2012 | 03/26/2012 | 07/03/2012 | 3002807907 | Penlon, Ltd. | Penlon, Ltd has identified the potential for the premature failure of a component in some Electronic Flowmeter sensor display unit PCB assemblies.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=185764 | Z-1181-2021 | Anesthesia gas-machine - Product Usage: intended for use in administering anesthesia while controlling the entire ventilation of patients with no ability to breathe, as well as in supporting patients with limited ability to breathe. | Flow-i Anesthesia System, Flow-c Anesthesia System, Flow-e Anesthesia System | 2 | 03/05/2021 |  | 09/28/2023 | 3025504317 | Getinge Group Logistics America, LLC | Maquet Critical Care AB received complaints where the FiCO2 value was not zero or as close to zero as expected and/or experienced System Check Out failures when docking the absorber to the Flow i Anesthesia machine.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=121944 | Z-1060-2014 | McKesson Anesthesia Care Product Usage: The McKesson Anesthesia Care system is a computer based system which collects, processes,and records data both through manual entry and from monitors which themselves are attached to patients, such as in the operating room environment. The system provides clinical decision support by communicating potential Adverse Drug Event alerts proactively during the pre-anesthesia evaluation and at the point-of-care. The system is generally indicated in the anesthetizing environment when the anesthesia provider decides to perform a patient assessment, to generate a paper and/or electronic record of the administration of anesthesia to a patient, and to document care. | McKesson Anesthesia Care System | 1 | 03/05/2014 | 03/05/2014 | 05/01/2014 | 3009662297 | MCKESSON TECHNOLOGIES INC. | There was an occurrence where the patient case data did not match the patient data when the case was recalled in the Anesthesia Care Record (ACR) in that it included data from another case.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161423 | Z-0847-2018 | A-Series A7 Anesthesia Delivery System; Model Number: 0632F-PA0000X (A7) Product Usage: The A-Series Anesthesia System is a device used to administer to a patient, continuously or intermittently, a general inhalation anesthetic, and to maintain a patient s ventilation. The A-Series is intended for use by licensed clinicians, for patients requiring anesthesia within a health care facility, and can be used for adult, pediatric, and infant populations. | A7 Anesthesia System | 2 | 03/02/2018 |  | 04/17/2019 | 2221819 | Mindray DS USA, Inc. dba Mindray North America | A software issue may result in the previous settings being applied instead of the default settings or the unit may skip the startup leak test.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161422 | Z-0846-2018 | A-Series A3/A5 Anesthesia Delivery System; Model Numbers: 0633F-01000-0X (A3) and 0631F-01000-0X (A5) Product Usage: The A-Series Anesthesia System is a device used to administer to a patient, continuously or intermittently, a general inhalation anesthetic, and to maintain a patient s ventilation. The A-Series is intended for use by licensed clinicians, for patients requiring anesthesia within a health care facility, and can be used for adult, pediatric, and infant populations. | A3/A5 Anesthesia System | 2 | 03/02/2018 |  | 04/17/2019 | 2221819 | Mindray DS USA, Inc. dba Mindray North America | A software issue may result in the previous settings being applied instead of the default settings or the unit may skip the startup leak test.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=31339 | Z-0545-04 | Datex-Ohmeda S/5 and AS/3 Anesthesia Delivery Units (The S/5 units are a later version of the AS/3 units. The conversion from AS3 to S5 occurred during the summer of 2000. This change was mainly a product name change, with a slight color variation to the shade of white being used externally on the machine. There were no design or feature and function modifications when the name change occurred.) | Datex-Ohmeda | 2 | 02/24/2004 | 07/20/2004 | 11/17/2005 | 2183066 | Datex-Ohmeda Inc | Anesthetic agents could be delivered at concentrations above the concentration set on the S/5 and AS/3 Anesthesia Delivery Units (ADU) having serial numbers within a specified range.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=143153 | Z-0804-2016 | Flow-i Anesthesia System Product Usage: The indication for the Flow-i Anesthesia System is administering inhalation anesthesia while controlling the entire ventilation of patients with no ability to breath, as well as in supporting patients with a limited ability to breath. The system is intended for use on neonatal to adult patient populations. The system is intended for use in hospital environments, except MRI environment by healthcare professionals trained in inhalation anesthesia administration. | Maquet Flow-i | 1 | 02/22/2016 | 02/22/2016 | 07/06/2016 | 3008355164 | Maquet Cardiovascular Us Sales, Llc | Electrical issues in the control, monitoring or panel sub-systems, caused by the printed circuit boards (PCBs) that control the sub-systems. The problems manifested as multiple re-starts of the sub-systems and, in a few instances, disabling of the system has occurred. Replacement PCBs are now available to correct the issues. Only the Model C20 is affected in the US.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=159630 | Z-0579-2018 | Carestation 620 A1, Carestation 650 A1, and Carestation 650c A1 | Carestation 620 A1, Carestation 650 A1, and Carestation 650c A1 | 2 | 02/09/2018 |  |  | 3003513129 | GE Medical Systems China Co., Ltd. | Possible incomplete seal between the CO2 absorber and the breathing circuit CO2 bypass port assembly of the Carestation 600 Series systems.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=159631 | Z-0580-2018 | Carestation 620 A2, Carestation 650 A2, and Carestation 650c A2 | Carestation 620 A2, Carestation 650 A2, and Carestation 650c A2 | 2 | 02/09/2018 |  |  | 3003513129 | GE Medical Systems China Co., Ltd. | Possible incomplete seal between the CO2 absorber and the breathing circuit CO2 bypass port assembly of the Carestation 600 Series systems.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211958 | Z-1074-2025 | Vaporizer Sevoflurane, Quik-Fil. Model Number: 6682285. Device is a component of Flow-i C20, Flow-i C30, Flow-i C40, Flow-c, Flow-e. For use during anesthesia. | Vaporizer Sevoflurane, Quik-Fil | 1 | 02/07/2025 | 02/07/2025 |  | 3013876692 | Getinge Usa Sales Inc | Sevoflurane used with the vaporizer may degrade to hydrogen fluoride, which could present a risk to patients or health care providers through inhalation and/or skin exposure. Potential harms include irritation of respiratory tract, which in worst case may lead to lung edema and/or severe hypocalcemia, blistering, superficial ulceration, and/or hypomagnesemia.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205337 | Z-0988-2024 | Flow-i C30 Anesthesia System, model 6677300 | Flow-i C30 | 2 | 02/06/2024 | 02/06/2024 |  | 3013876692 | Getinge Usa Sales Inc | Potential cybersecurity vulnerability that could lead to a Denial of Service (DoS) attack, tampering, or remote code execution via remote login to the Flow Anesthesia System, any of which could lead to a hazardous situation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205339 | Z-0990-2024 | Flow-c Anesthesia System, model 6887700 | Flow-c | 2 | 02/06/2024 | 02/06/2024 |  | 3013876692 | Getinge Usa Sales Inc | Potential cybersecurity vulnerability that could lead to a Denial of Service (DoS) attack, tampering, or remote code execution via remote login to the Flow Anesthesia System, any of which could lead to a hazardous situation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205336 | Z-0987-2024 | Flow-i C20 Anesthesia System, model 6677200 | Flow-i C20 | 2 | 02/06/2024 | 02/06/2024 |  | 3013876692 | Getinge Usa Sales Inc | Potential cybersecurity vulnerability that could lead to a Denial of Service (DoS) attack, tampering, or remote code execution via remote login to the Flow Anesthesia System, any of which could lead to a hazardous situation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205340 | Z-0991-2024 | Flow-e Anesthesia System, model 6887900 | Flow-e | 2 | 02/06/2024 | 02/06/2024 |  | 3013876692 | Getinge Usa Sales Inc | Potential cybersecurity vulnerability that could lead to a Denial of Service (DoS) attack, tampering, or remote code execution via remote login to the Flow Anesthesia System, any of which could lead to a hazardous situation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=205338 | Z-0989-2024 | Flow-i C40 Anesthesia System, model 6677400 | Flow-i C40 | 2 | 02/06/2024 | 02/06/2024 |  | 3013876692 | Getinge Usa Sales Inc | Potential cybersecurity vulnerability that could lead to a Denial of Service (DoS) attack, tampering, or remote code execution via remote login to the Flow Anesthesia System, any of which could lead to a hazardous situation.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=204697 | Z-0971-2024 | Atlan, Gas Machine for Anesthesia or Analgesia, Catalog numbers 8621500 (Atlan A350) and 8621600 (Atlan A350 XL) | Atlan | 2 | 02/01/2024 | 02/01/2024 |  | 3001104093 | Draegerwerk Ag & Co. Kgaa | Internal backup battery may fail spontaneously while the Atlan is being operated without AC mains supply, resulting in unexpected shutdown of the device while running on batteries.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197515 | Z-0961-2023 | Getinge Flow-e Anesthesia System, Part No. 6887900 | Getinge Flow-e | 1 | 01/26/2023 | 01/26/2023 |  | 3013876692 | Getinge Usa Sales Inc | Due to a software bug, under certain conditions, pressure cannot be built up resulting in no ventilation. If gas delivery is stopped, a sustained decrease in delivered O2 concentration may lead to hypoventilation and hypoxia.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197513 | Z-0959-2023 | Getinge Flow-i Anesthesia Systems Flow-i C20, Part no. 6677200 Flow-i C30, Part no. 6677300 Flow-i C40, Part no. 6677400 | Getinge Flow-i | 1 | 01/26/2023 | 01/26/2023 |  | 3013876692 | Getinge Usa Sales Inc | Due to a software bug, under certain conditions, pressure cannot be built up resulting in no ventilation. If gas delivery is stopped, a sustained decrease in delivered O2 concentration may lead to hypoventilation and hypoxia.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=197514 | Z-0960-2023 | Getinge Flow-c Anesthesia System, Part No. 6887700 | Getinge Flow-c | 1 | 01/26/2023 | 01/26/2023 |  | 3013876692 | Getinge Usa Sales Inc | Due to a software bug, under certain conditions, pressure cannot be built up resulting in no ventilation. If gas delivery is stopped, a sustained decrease in delivered O2 concentration may lead to hypoventilation and hypoxia.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178323 | Z-0814-2020 | Carestation 620/650/650c A2 Product Usage: anesthesia systems are intended to provide general inhalation anesthesia and ventilatory support to a wide range of patients (pediatric and adult. | Carestation 620/650/650c A2 | 1 | 01/23/2020 | 01/23/2020 | 11/30/2023 | 2126677 | GE Healthcare, LLC | GE Healthcare has become aware that there is a potential for a loose cable connection inside specific manufactured anesthesia devices. This would cause a loss of mechanical ventilation and the system will provide high priority audio and visual alarms. Loss of mechanical ventilation could lead to hypoxia if the clinician does not intervene.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178322 | Z-0813-2020 | Carestation 620/650/650c A1 Product Usage: anesthesia systems are intended to provide general inhalation anesthesia and ventilatory support to a wide range of patients (pediatric and adult. | Carestation 620/650/650c A1 | 1 | 01/23/2020 | 01/23/2020 | 11/30/2023 | 2126677 | GE Healthcare, LLC | GE Healthcare has become aware that there is a potential for a loose cable connection inside specific manufactured anesthesia devices. This would cause a loss of mechanical ventilation and the system will provide high priority audio and visual alarms. Loss of mechanical ventilation could lead to hypoxia if the clinician does not intervene.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142562 | Z-0681-2016 | GE Healthcare, Aisys CS2,1011-9050-000; Avance Amingo,1009-9002-000, & Avance CS2, 1009-9050-000. The GE Datex-Ohmeda Aisys Anesthesia System is intended to provide general inhalation anesthesia and ventilatory support to a wide range of patients (neonatal, pediatric, adult). The device is intended for volume or pressure control ventilation. The Aisys is not suitable for use in a MRI environment. | GE Healthcare, Aisys CS2; Avance Amingo; Avance CS2 | 2 | 01/21/2016 | 01/21/2016 | 09/21/2016 | 2126677 | GE Medical Systems, LLC | The Drive Gas Check Valve of certain Anesthesia products could become stuck in a fixed open position, which could cause pressure to build in the mechanical ventilation cycle. If this issue is left unresolved, it could result in excessive or prolonged pressure in the patient breathing circuit during ventilation potentially resulting in barotrauma.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142559 | Z-0679-2016 | GE Healthcare, Aestiva MRI, 1006-9310-000. This version of the Datex-Ohmeda 7900 ventilator is used in Datex-Ohmeda Aestiva/5 anesthesia systems. It is a microprocessor based, electronically controlled, pneumatically driven ventilator that provides patient ventilation during surgical procedures. The Aestiva/5 MRI anesthesia system provides the functional feature set offered by the conventional Aestiva/5 to the clinician with the added ability to be used in the MR environment. | Aestiva MRI | 2 | 01/21/2016 | 01/21/2016 | 09/21/2016 | 2126677 | GE Medical Systems, LLC | The Drive Gas Check Valve of certain Anesthesia products could become stuck in a fixed open position, which could cause pressure to build in the mechanical ventilation cycle. If this issue is left unresolved, it could result in excessive or prolonged pressure in the patient breathing circuit during ventilation potentially resulting in barotrauma.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142568 | Z-0682-2016 | Anesthesia device service kits. ASSY-MSN, VENT ENGINE ASSY, 1503-8101-000 ASSY-MSN, VENT ENGINE ASSY 7900, 1009-8216-000 ASSY-MSN, VENT ENGINE ASSY AESPIRE 7900 1009-8423-000 ASSY-MSN, VENT ENGINE ASSY, AESTIVA 7900 MRI 1503-8102-000 ASSY-MSN, VLV ASSY DR GAS CHK BCG, Manufacturing assembly - Buy 1503-3006-000 Service kits containing the ventilator Drive Gas Check Valve shipped from April 20, 2015 through October 2015 are also affected and could have been installed on Aespire 7900, Aespire View, Aestiva 7900, Aestiva MRI, Aisys, Aisys CS2, Avance, Avance CS2 Anesthesia, and Amingo devices. (Service kit part numbers: 1009-8216-000, 1503-3006-000, 1503-8102-000, 1009-8423-000, 1503-8101-000). | Aespire, Aestiva, Aisys, Avance, Anesthesia device service kits | 2 | 01/21/2016 | 01/21/2016 | 09/21/2016 | 2126677 | GE Medical Systems, LLC | The Drive Gas Check Valve of certain Anesthesia products could become stuck in a fixed open position, which could cause pressure to build in the mechanical ventilation cycle. If this issue is left unresolved, it could result in excessive or prolonged pressure in the patient breathing circuit during ventilation potentially resulting in barotrauma.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142560 | Z-0680-2016 | GE Healthcare, Aisys, 1011-9000-000. The GE Datex-Ohmeda Aisys Anesthesia System is intended to provide general inhalation anesthesia and ventilatory support to a wide range of patients (neonatal, pediatric, adult). The device is intended for volume or pressure control ventilation. The Aisys is not suitable for use in a MRI environment. | Aisys | 2 | 01/21/2016 | 01/21/2016 | 09/21/2016 | 2126677 | GE Medical Systems, LLC | The Drive Gas Check Valve of certain Anesthesia products could become stuck in a fixed open position, which could cause pressure to build in the mechanical ventilation cycle. If this issue is left unresolved, it could result in excessive or prolonged pressure in the patient breathing circuit during ventilation potentially resulting in barotrauma.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142557 | Z-0677-2016 | GE Healthcare, Aespire 7900, 1009-9012-000 & Aespire View, 1009-9212-000. The family of GE Datex-Ohmeda Aespire anesthesia systems with 7900 ventilator (Aespire 7900 and Aespire View) is intended to provide general inhalation anesthesia and ventilatory support to a wide range of patients. The devices are intended for volume or pressure control ventilation. | Aespire 7900; Aespire View | 2 | 01/21/2016 | 01/21/2016 | 09/21/2016 | 2126677 | GE Medical Systems, LLC | The Drive Gas Check Valve of certain Anesthesia products could become stuck in a fixed open position, which could cause pressure to build in the mechanical ventilation cycle. If this issue is left unresolved, it could result in excessive or prolonged pressure in the patient breathing circuit during ventilation potentially resulting in barotrauma.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190921 | Z-0501-2022 | Draeger Fabius MRI Anesthesia Machine, Catalog Number 8607300 | Fabius MRI | 2 | 01/14/2022 |  | 04/17/2023 | 2517967 | Draeger Medical, Inc. | 26 Fabius MRI Anesthesia machines were distributed with internal hoses routing Oxygen to the Oxygen Flush and Oxygen Flow Control interchanged. If the hospital O2 Central supply pressure is not stable, the O2 fresh gas flow could deviate from the adjusted value. If the deviation is not noticed, insufficient oxygenation of the patient may occur.
2 | BSZ | GAS-MACHINE, ANESTHESIA | 868.5160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=190581 | Z-0468-2022 | Vaporizer Sevoflurane Maquet Filling, Part No. 6886601 Component of the following systems: Product Code(s) UPN/UDI: Flow-c Anesthesia System 6887700 07325710009765 Flow-e Anesthesia System 6887900 07325710010457 Flow-i Anesthesia System C20 6888520 07325710010617 Flow-i Anesthesia System C30 6888530 07325710010624 Flow-i Anesthesia System C40 6888540 07325710010631 Flow-i C20 Anesthesia System 6677200 07325710001349 Flow-i C30 Anesthesia System 6677300 07325710001349 Flow-i C40 Anesthesia System 6677400 07325710001318 | Vaporizer Sevoflurane, Maquet Filling | 1 | 01/10/2022 |  |  | 3013876692 | Getinge Usa Sales Inc | A potential chemical degradation of Sevoflurane by Lewis acids (metal oxides and metal halides) has been discovered that may result in inhalation and/or skin exposure of hydrogen fluoride.
2
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=112681 | Z-0606-2013 | GE Giraffe and Panda T-Piece Resuscitation System labeled ***T-Piece Neonatal Patient Circuit Kit Disposable with Size 1 Mask***M1091335 QTY: 10 T-Piece Disposable Circuit, QTY: 10, M1091316 T-Piece Disposable Circuit Kit with Mask Size O; M1091365 QTY: 10 T-Piece Disposable Circuit Kit with Mask, Size 1, Lot 1981713 *** Ohmeda Medical, A Division of Datex-Ohmeda, Inc. A General Electric Company Laurel, MD 20723 USA MADE IN USA*** Provides the basic equipment required for pulmonary resuscitation of infants. | Giraffe and Panda Warmer T-Piece Resuscitation System | 1 | 12/26/2012 | 12/26/2012 | 04/14/2014 | 2126677 | GE Healthcare, LLC | Disposable T-Piece circuits used with resuscitation systems are unable to achieve maximum Peak Inspiratory Pressure (PIP) during the pre-use check out procedure.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189265 | Z-0163-2022 | Pneupac paraPac plus 300 ventilator kit, REF P300NGB | Pneupac paraPac plus 300 ventilator kit | 3 | 10/22/2021 |  | 08/11/2023 | 3012307300 | Smiths Medical ASD Inc. | Two devices were labeled with the same serial number
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209941 | Z-0004-2025 | Neo-Tee T-Piece Resuscitator. Part Numbers 1050805, 1050832. For pediatric emergency respiratory support. | Neo-Tee T-Piece Resuscitator | 1 | 10/02/2024 | 10/02/2024 |  | 1024404 | Mercury Enterprises, Inc. dba Mercury Medical | Manual resuscitator circuit flow controller may come apart due to a missing manufacturing step.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=208998 | Z-3090-2024 | paraPAC plus 310 ventilator kit with internal PEEP and CPAP, Item Number P310N | Smiths Medical | 1 | 09/20/2024 | 09/20/2024 |  | 3012307300 | Smiths Medical ASD Inc. | Smiths Medical became aware of an issue related to a potential for the patient outlet connector to loosen/detach from the paraPAC Plus P300 and P310 ventilators impacting the active ventilation function.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209599 | Z-3091-2024 | paraPAC 300 ventilator kit without internal PEEP and CPAP, Item Number P300N | Smiths Medical | 1 | 09/20/2024 | 09/20/2024 |  | 3012307300 | Smiths Medical ASD Inc. | Smiths Medical became aware of an issue related to a potential for the patient outlet connector to loosen/detach from the paraPAC Plus P300 and P310 ventilators impacting the active ventilation function.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209122 | Z-2973-2024 | Pneupac paraPAC plus 300 ventilator kit, REF P300NXX | PneuPAC | 1 | 09/11/2024 | 09/11/2024 |  | 3012307300 | Smiths Medical ASD Inc. | There is a potential for inadvertent tidal volume knob movement from the original setting when set at high (1000-1500 mL) and low (70-150 mL) settings.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=209123 | Z-2974-2024 | Pneupac paraPAC plus 310 ventilator kit with PEEP and CPAP, REF P310NXX | Pneupac | 1 | 09/11/2024 | 09/11/2024 |  | 3012307300 | Smiths Medical ASD Inc. | There is a potential for inadvertent tidal volume knob movement from the original setting when set at high (1000-1500 mL) and low (70-150 mL) settings.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=182621 | Z-2805-2020 | PVA, PREVENT is an FDA-authorized emergency ventilator designed to treat patients suffering from COVID-19 respiratory failure. This motorized device automatically compresses a resuscitator bag to provide air to the lungs of a COVID-19 patient, eliminating the human error and fatigue that results from manual actuation. PREVENT is indicated for emergency resuscitation with appropriate patient monitoring on adult patients that require mechanical respiratory support. | PVA, PREVENT | 2 | 08/14/2020 |  |  | 3001151872 | Precision Valve & Automation,Inc | While operating the machine in "Run" mode an unexcepted event may occur during the error-checking for stepper movements and pressure values. The result of this event is a false alarm being thrown and in rare cases results in an incomplete move which will affect the current inhale/exhale cycle.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=52551 | Z-0991-2007 | Pneupac VR1 Standard and VR1 Responder is a hand held portable, time cycled, gas powered, flow generator ventilatory resuscitators that are intended for emergency resuscitation by medical personnel, paramedics and ambulance technicians inside and outside hospital and for ventilation by medical personnel inside and outside hospital in emergency situations and for intra- and inter- hospital transport. | Pneupac VR1 | 2 | 07/03/2007 | 07/04/2007 | 10/10/2008 | 2183502 | Smiths Medical PM, Inc. | Smiths Medical is conducting a recall of 93 Pneupac VR1 Emergency & Transport Ventilators following the identification of a problem which may exist with the demand valve that could cause the ventilator to stop cycling & to deliver continuous flow output to the patient.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=154926 | Z-2092-2017 | VORTRAN Automatic Resuscitator, VARPlus with Extension Kit; Model PCE-5012; Qty: 10 each. Anesthesiology: This device is to be used by properly trained personnel to deliver emergency, short term, constant flow, pressure cycled ventilatory support. | VORTRAN Automatic Resuscitator, VAR-Plus | 2 | 05/13/2017 | 04/25/2017 | 10/17/2017 | 1000307109 | Vortran Medical Technology 1, Inc | Vortran received a customer complaint that the spin-nut-DSS connector was assembled in the reverse direction on the oxygen tubing; making the connector unable to be screwed on due to the wrong thread direction.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=154925 | Z-2091-2017 | VORTRAN Automatic Resuscitator, VARPlus; Model PTE-5002; Qty: 10 each Anesthesiology: This device is to be used by properly trained personnel to deliver emergency, short term, constant flow, pressure cycled ventilatory support. | VORTRAN Automatic Resuscitator | 2 | 05/13/2017 | 04/25/2017 | 10/17/2017 | 1000307109 | Vortran Medical Technology 1, Inc | Vortran received a customer complaint that the spin-nut-DSS connector was assembled in the reverse direction on the oxygen tubing; making the connector unable to be screwed on due to the wrong thread direction.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=116501 | Z-1125-2013 | Instrumentation Industries BE183-SU and BE 183 SU-12 Manual Jet Ventilators; BE 183-SUR and BE 183 SUR-12 Manual Jet Ventilators with Regulator and Gauge; BE 183-2 4 ft. Small Bore Tubing Assembly w/Inline Filter; and BE 183-R Regulator The BE 183-SUR and BE 183-SU Manual Jet Ventilators are designed to provide transtracheal ventilation in specific emergency situations of upper airway obstructions. | Manual Jet Ventilators, Manual Jet Ventilators w/Regulator and Guage, Small Bore Tubing Assembly | 2 | 04/16/2013 | 04/16/2013 | 07/05/2013 | 2518436 | Instrumentation Industries Inc | Firm initiated an update to Directions for Use necessitating replacement of previously distributed Directions for Use.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=126050 | Z-1440-2014 | Draeger Oxylog 3000 plus ventilators, emergency and transport ventilator. Part number: 5704833 | Draeger Oxylog 3000 plus ventilators | 1 | 04/10/2014 | 04/10/2014 | 10/31/2014 | 2510954 | Draeger Medical Systems, Inc. | The potentiometers (knobs) for setting the ventilation parameters may have contact interrupted and the device discontinues ventilation, gives an audible alarm and displays the error message Poti unplugged.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171092 | Z-0976-2019 | e500 Automatic Transport Ventilator, Model Number 01EVE500 | e500 Automatic Transport Ventilator | 1 | 04/04/2019 | 04/04/2019 | 06/29/2023 | 3002807836 | O-Two Medical Technologies, Inc. | There is a potential for the unit to catch fire due to one of the screen retaining screws being incorrectly inserted. Due to this improper screw insertion the screw sat higher than normal allowing it to contact the underside of the main board causing the screw to rub against the bottom side of the board, wearing away the insulation layer and creating a short circuit.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171094 | Z-0978-2019 | e700 Automatic Transport Ventilator, Model Number 01EVE700 | o_two e700 Automatic Transport Ventilator | 1 | 04/04/2019 | 04/04/2019 | 06/29/2023 | 3002807836 | O-Two Medical Technologies, Inc. | There is a potential for the unit to catch fire due to one of the screen retaining screws being incorrectly inserted. Due to this improper screw insertion the screw sat higher than normal allowing it to contact the underside of the main board causing the screw to rub against the bottom side of the board, wearing away the insulation layer and creating a short circuit.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171093 | Z-0977-2019 | e600 Automatic Transport Ventilator, Model Number 01EVE600 | o_two e600 Automatic Transport Ventilator | 1 | 04/04/2019 | 04/04/2019 | 06/29/2023 | 3002807836 | O-Two Medical Technologies, Inc. | There is a potential for the unit to catch fire due to one of the screen retaining screws being incorrectly inserted. Due to this improper screw insertion the screw sat higher than normal allowing it to contact the underside of the main board causing the screw to rub against the bottom side of the board, wearing away the insulation layer and creating a short circuit.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=206298 | Z-1274-2024 | smiths medical Pneupac paraPAC plus 310 ventilator kit with PEEP and CPAP, REF P310NUS | smiths medical Pneupac | 1 | 03/19/2024 | 03/19/2024 |  | 3012307300 | Smiths Medical ASD Inc. | When a paraPAC plus ventilator is switched to the operating mode of VENTILATE, the ventilator may intermittently provide continuous positive gas flow instead of the intended cycling like a human breath. This non-cycling and continuous positive gas flow when in the cycling mode, is a malfunction, not allowing the ventilator to properly function as designed.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=206299 | Z-1275-2024 | smiths medical Pneupac paraPAC plus 300 ventilator kit without PEEP and CPAP, REF P300NUS | smiths medical Pneupac | 1 | 03/19/2024 | 03/19/2024 |  | 3012307300 | Smiths Medical ASD Inc. | When a paraPAC plus ventilator is switched to the operating mode of VENTILATE, the ventilator may intermittently provide continuous positive gas flow instead of the intended cycling like a human breath. This non-cycling and continuous positive gas flow when in the cycling mode, is a malfunction, not allowing the ventilator to properly function as designed.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=124480 | Z-1070-2014 | GE Healthcare - Giraffe Warmer;Panda Warmer;Panda iRes Warmer, Giraffe Stand-alone Infant Resuscitation System. Product Usage: The T-Piece Resuscitation System [ventilator, emergency powered resuscitator]/Panda Bag and Mask Resuscitation System provides the basic equipment required for pulmonary resuscitation of infants. Pulmonary resuscitation includes practices necessary to establish a clear airway and provide oxygen or air/oxygen mixtures and/or manual ventilation to the infant. These are clinical practices that represent the established standard of care. | Giraffe and Panda Warmer T-Piece Resuscitation System | 1 | 02/25/2014 | 02/25/2014 | 07/22/2016 | 2126677 | GE Healthcare, LLC | Oxygen and air wall inlet fittings on the back panel of the resuscitation systems were reversed during assembly.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=125259 | Z-1072-2014 | GE Healthcare - Giraffe and Panda iRes Infant Warmer Product Usage: The infant radiant warmers provide infrared heat in a controlled manner to neonates who are unable to thermo-regulate based on their own physiology. Infant radiant warmers may be used to facilitate the neonates transition to the external environment or to provide a controlled open environment. An optional integrated Sp02 monitoring feature may be used for continuous noninvasive monitoring of functional oxygen saturation of arterial hemoglobin and pulse rate. | Emergency Powered Ventilator Resusicitator | 1 | 02/25/2014 | 02/25/2014 | 07/22/2016 | 2126677 | GE Healthcare, LLC | Oxygen and air wall inlet fittings on the back panel of the resuscitation systems were reversed during assembly.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=114723 | Z-0748-2013 | GE Healthcare - Giraffe and Panda T-Piece Resuscitation System The T-piece Resuscitation System [ventilator, emergency powered resuscitator]/Panda Bag and Mask Resuscitation System provides the basic equipment required for pulmonary resuscitation of infants. | Giraffe and Panda Warmer T-Piece Resuscitation System | 1 | 02/08/2013 | 02/08/2013 | 01/18/2017 | 2126677 | GE Healthcare, LLC | Oxygen and air wall inlet fittings on the back panel of the resuscitation systems were reversed during assembly.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=125032 | Z-0916-2014 | Stand-Alone Resuscitation, T-Piece, part numbers 2063771-001, 2063772-001, 2063773-001, 2063774-001, 2063775-001, M1226442, M1241420, M1241421, M1241424, M1241425, M1241427, M1241430, M1241461, M1241463, M1241464, and M1241589; labeled in part ***GE Healthcare 8880 Gorman Road Laurel, MD 20723 ***www.gehealthcare.com*** Provides the basic equipment required for pulmonary resuscitation of infants. | Stand-Alone Resuscitation, T-Piece | 2 | 02/03/2014 | 02/03/2014 | 06/01/2015 | 1121732 | GE Healthcare | Medical device adjustment knob may be loose and affect the accuracy of blended Air/Oxygen gas mixture to patients.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=125033 | Z-0917-2014 | Upgrade Kit, Resuscitation, T-Piece, part numbers M1139417, M1192226 and M1223628, labeled in part ***GE Healthcare 8880 Gorman Road Laurel, MD 20723 ***www.gehealthcare.com*** Provides the basic equipment required for pulmonary resuscitation of infants. | Upgrade Kit, Resuscitation, T-Piece | 2 | 02/03/2014 | 02/03/2014 | 06/01/2015 | 1121732 | GE Healthcare | Medical device adjustment knob may be loose and affect the accuracy of blended Air/Oxygen gas mixture to patients.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=106467 | Z-0870-2012 | Model 73X Uni-Vent Product Usage: Intended for use at mass casualty incidents. | UNI-VENT MODEL 73X | 2 | 01/25/2012 |  | 05/23/2013 | 1220908 | Impact Instrumentation, Inc. | A piece of internal tubing in the 73X portable ventilator is of incorrect specifications for the device.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=43414 | Z-0436-06 | Smiths Pneupac babyPAC portable ventilator, model 100. Designed for use by qualified medical caregivers, paramedics and other trained personnel, for hospital, emergency and transport ventilation of patients during respiratory distress or insufficiency. | Smiths Pneupac babyPAC 100 | 2 | 01/23/2006 | 01/24/2006 | 12/26/2006 | 2183502 | Smiths Medical PM, Inc. | A potential safety related problem associated with the Pnuepac babyPac Ventilator model 100 has been discovered in that the interface between the air input hose and the air input fitting may leak, and if tightened too much, the air input fitting may loosen when trying to remove the air hose.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=141479 | Z-0639-2016 | Mercury Medical Neo-Tee T-Piece Resuscitator and T-Piece Circuit. The Neo-Tee T-Piece Resuscitator is a gas powered emergency resuscitator intended to provide emergency respiratory support by means of a face mask or a tube inserted into a patient's airway. It is intended for use with neonates and infants weighing less than 10kg (22 lb). The T-piece resuscitator circuit is a component used in conjunction with the T-piece resuscitator devices | Mercury Medical Neo-Tee | 2 | 01/12/2016 |  | 10/13/2017 | 1024404 | Mercury Enterprises, Inc. dba Mercury Medical | Incorrect use of T-Piece resuscitators will result in complete occlusion of the patient circuit and consequent failure to ventilate the patient.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142381 | Z-0586-2016 | MOVES System, an emergency ventilator with suction, oxygen concentrator and multi-parameter patient monitoring capabilities. The MOVES System is a portable computer controlled electrically powered transport emergency ventilator intended to provide continuous or intermittent ventilatory support for the care of adults who require ventilation. | Portable Life Support System | 1 | 01/11/2016 | 01/11/2016 | 04/12/2019 | 3005803793 | Thornhill Research Inc | Batteries found unresponsive after storage, not recognized by main device, and are perpetually enabled.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=142382 | Z-0587-2016 | MOVES System Lithium Polymer Batteries The MOVES System is a portable computer controlled electrically powered transport emergency ventilator intended to provide continuous or intermittent ventilatory support for the care of adults who require ventilation. | Portable Life Support System | 1 | 01/11/2016 | 01/11/2016 | 04/12/2019 | 3005803793 | Thornhill Research Inc | Batteries found unresponsive after storage, not recognized by main device, and are perpetually enabled.
2 | BTL | VENTILATOR, EMERGENCY, POWERED (RESUSCITATOR) | 868.5925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211660 | Z-0821-2025 | Neo-Tee T-Piece Resuscitator, for pediatric emergency respiratory support, Part: 1050805, 1050808, 1050809, 1050810, 1050811, 1050814, 1050832, 1050839, 1050840, 1050841, 1050842 | Neo-Tee T-Piece Resuscitator | 1 | 01/10/2025 | 01/10/2025 |  | 1024404 | Mercury Enterprises, Inc. dba Mercury Medical | Gas powered emergency resuscitator has a circuit flow controller manufactured with an undersized spring, which may cause resuscitator not reach the high and low ends of the PIP and PEEP pressure ranges, which may cause a loss of positive pressure, impacting effective ventilation of the patient.
2
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=204714 | Z-0621-2024 | Atrium Oasis Drain Single W / AC, Part No. 3600-100 | Atrium Oasis Drain Single W / AC | 2 | 12/28/2023 | 12/28/2023 |  | 3011175548 | Atrium Medical Corporation | A sterilization nonconformance prematurely aged the product by 48.3 days, resulting in the expiration date on the label not being accurate. This nonconforming product was inadvertently distributed to customers. Increased risk to patients and healthcare providers using the product after the shortened expiration date of November 17, 2025.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=30408 | Z-0227-04 | Autovac TL. Autotransfusion System for Orthopedic Wound Drainage. The product is shipped in cartons containing 6 units | Autovac TL | 3 | 12/16/2003 | 12/17/2003 | 01/20/2004 | 2518417 | Boehringer Laboratories Inc | Inadequate bond where the tube attaches to the inlet of the bulb.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=65455 | Z-0358-2008 | Haemonetics cardioPAT System, Product List Number 02050, Haemonetics Corporation, Braintree, MA 02184,(Cardiovascular Perioperative Autotransfusion) | Haemonetics cardioPAT System | 2 | 12/07/2007 | 12/08/2007 | 03/19/2012 | 1219343 | Haemonetics Corporation | Labeling (additional) use instructions: to clarify proper position of cardioPAT device and post-op line to prevent fluid build up from patient's chest
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=204036 | Z-0486-2024 | Atrium Oasis Dry Suction Water Seal Chest Drain, Part Number: 3650-100, DRAIN, OASIS BRU W/AC, 3650 ATS BLOOD RECOVERY | Atrium Oasis Dry Suction Water Seal Chest Drain | 2 | 12/01/2023 | 12/02/2023 |  | 3011175548 | Atrium Medical Corporation | Product was re-processed and re-sterilized by a third party using packaging configurations and a sterilization process which are not approved by the manufacture and lack validation.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=40930 | Z-0174-06 | cardioPAT Cardiovascular Perioperative Autotransfusion System, Model: LN 2050-110-E and Disposables: cardioPAT Replacement RBC Bags, LN 00270 cardioPAT Replacement Waste Bags, LN 00271 cardioPAT Replacement Intraoperative Suction Set, LN 00272 cardioPAT Replacement Postoperative Suction Set, LN 00273 cardioPAT cardiovascular Perioperative Autotransfusion System Disposable Set, LN 00274 cardioPAT Replacement One-Way Valve Assembly LN 00275 | Haemonetics | 2 | 11/22/2005 | 11/23/2005 | 05/24/2012 | 1219343 | Haemonetics Corporation | Blood clotting of patient post operative while on the cardioPAT Cardiovascular Perioperative Autotransfusion System
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203545 | Z-0318-2024 | Express Dry Seal Chest Drains (Express): 4000-100N DRAIN, EXPRESS, SINGLE; 4050-100N DRAIN, EXPRESS, BRU | Express Dry Seal Chest Drains (Express) | 2 | 11/16/2023 | 11/16/2023 |  | 3011175548 | Atrium Medical Corporation | Instructions for Use (IFU) for the Atrium Ocean, Oasis, and Express chest drains do not provide sufficient precaution instruction for proper set up of catheter(s) and patient tube connections with single collection chamber chest drains
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203544 | Z-0317-2024 | Oasis Dry Suction Water Seal Chest Drains (Oasis): (1)3600-100 DRAIN, OASIS SINGLE W/AC; (2)3612-100 DRAIN, OASIS PEDI A/C; (3)3650-100 DRAIN, OASIS BRU W/AC; | Oasis Dry Suction Water Seal Chest Drains (Oasis) | 2 | 11/16/2023 | 11/16/2023 |  | 3011175548 | Atrium Medical Corporation | Instructions for Use (IFU) for the Atrium Ocean, Oasis, and Express chest drains do not provide sufficient precaution instruction for proper set up of catheter(s) and patient tube connections with single collection chamber chest drains
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203543 | Z-0316-2024 | Ocean Water Seal Chest Drains (Ocean): Products REF Number Product Name (1)2002-000 DRAIN, OCEAN SINGLE W/AC,S; (2)2002-040 DRAIN, OCEAN SINGLE,PEDI CONNECTOR; (3)2002-100 DRAIN, OCEAN SINGLE W/AC; (4)2002-300 DRAIN, OCEAN SINGLE W/S; (5) 2002-400 DRAIN, OCEAN SINGLE ; (6)2012-320 DRAIN, OCEAN PEDI W/S; (7)2050-000 DRAIN, OCEAN BRU W/AC,S | Ocean Water Seal Chest Drains (Ocean) | 2 | 11/16/2023 | 11/16/2023 |  | 3011175548 | Atrium Medical Corporation | Instructions for Use (IFU) for the Atrium Ocean, Oasis, and Express chest drains do not provide sufficient precaution instruction for proper set up of catheter(s) and patient tube connections with single collection chamber chest drains
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94949 | Z-0341-2011 | The box label states: "ATF 40 Fast Start Kit 9108491...STERILE...FRESENIUS KABI...Fresenius Kabi AG 61345 Bad Hamburg Germany...Federal (USA) Law restricts this device to sale by or on order of a physician". The ATF 40 Fast Start Kit contains one each AT1 Autotransfusion Set, ATS Suction Line, and the ATR40 Autotransfusion Reservoir. ATS is a suction line to aspirate and add anticoagulate to shed blood from the surgical field into an autotransfusion reservoir (ATR). The ATS has a inch. suction port and is connected to a vacuum source. ATR is an autotransfusion reservoir to collect, de-foam, filter, and store blood prior to processing. The ATR reservoirs are offered with either a 120 micron or a 40 micron filter. ATF 40 Autotransfusion set is a disposable set, dedicated component of the Fresenius CATS Continuous Autotransfusion System, which in turn is an Autotransfusion device for the processing of autologous shed blood collected intraoperatively and postoperatively to obtain washed packed red blood cells for reinfusion. | ATF 40 (item # 9108491) | 2 | 11/10/2010 | 11/10/2010 | 06/15/2011 | 3007344667 | Fresenius Kabi, LLC | Using AT1 Autotransfusion Set from specific batches in combination with the auto transfusion device CATS Continuous Autotransfusion has led to insufficient suction of shed blood from the collection reservoir. The results in the failure message "alarm blood flow" and the CATS device stops in a secure mode.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94954 | Z-0342-2011 | The box label states: "ATF 120 Fast Start Kit 9108501...STERILE...FRESENIUS KABI...Fresenius Kabi AG 61345 Bad Hamburg Germany...Federal (USA) Law restricts this device to sale by or on order of a physician". The ATF120 Fast Start Kit contains one each AT1 Autotransfusion Set, ATS Suction Line, and the ATR120 Autotransfusion Reservoir. ATS is a suction line to aspirate and add anticoagulate to shed blood from the surgical field into an autotransfusion reservoir (ATR). The ATS has a 1/4 inch. suction port and is connected to a vacuum source. ATR is an autotransfusion reservoir to collect, de-foam, filter, and store blood prior to processing. The ATR reservoirs are offered with either a 120 micron or a 40 micron filter. ATF 120 Autotransfusion set is a disposable set, dedicated component of the Fresenius CATS Continuous Autotransfusion System, which in turn is an Autotransfusion device for the processing of autologous shed blood collected intraoperatively and postoperatively to obtain washed packed red blood cells for reinfusion. | ATF 120 (item # 9108501) | 2 | 11/10/2010 | 11/10/2010 | 06/15/2011 | 3007344667 | Fresenius Kabi, LLC | Using AT1 Autotransfusion Set from specific batches in combination with the auto transfusion device CATS Continuous Autotransfusion has led to insufficient suction of shed blood from the collection reservoir. The results in the failure message "alarm blood flow" and the CATS device stops in a secure mode.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94446 | Z-0340-2011 | The box label states: "AT1 Autotransfusion Set 9005101 STERILE...FRESENIUS KABI...Fresenius Kabi AG 61345 Bad Hamburg Germany...Federal (USA) Law restricts this device to sale by or on order of a physician". The AT1 Autotransfusion Set is distributed by Fresenius Kabi on its own or as part of a ATF40 kit or ATF120 kit. AT1 Autotransfusion set is a disposable set, dedicated component of the Fresenius CATS Continuous Autotransfusion System, which in turn is an Autotransfusion device for the processing of autologous shed blood collected intraoperatively and postoperatively to obtain washed packed red blood cells for reinfusion. | AT1 Autotransfusion Set (item #9005101) | 2 | 11/10/2010 | 11/10/2010 | 06/15/2011 | 3007344667 | Fresenius Kabi, LLC | Using AT1 Autotransfusion Set from specific batches in combination with the auto transfusion device CATS Continuous Autotransfusion has led to insufficient suction of shed blood from the collection reservoir. The results in the failure message "alarm blood flow" and the CATS device stops in a secure mode.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=69077 | Z-1480-2008 | Zimmer H.A.S. Autotransfusion System, H.A.S. Kit with 1/8" Drain , 3,2mm, 10Fr, PVC, 107cm, medium, sterile, latex free, Rx only; Zimmer, Swindon, Wiltshire, United Kingdom; Zimmer, Dover, Ohio; REF 00-2555-020-05. | Zimmer H.A.S. Autotransfusion System H.A.S. Kit with 1/8" Drain | 2 | 09/19/2008 | 09/19/2008 | 11/12/2009 | 1000220733 | Zimmer Inc. | The sterility of the device may be compromised due to the possible presence of slits or pinholes in the packaging.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=69096 | Z-1483-2008 | Zimmer H.A.S. Autotransfusion System, H.A.S. Replacement Kit, sterile, latex free, Rx only; Zimmer, Swindon, Wiltshire, United Kingdom; Zimmer, Dover, Ohio; REF 00-2555-050-05. | Zimmer H.A.S. Autotransfusion System Replacement Kit | 2 | 09/19/2008 | 09/19/2008 | 11/12/2009 | 1000220733 | Zimmer Inc. | The sterility of the device may be compromised due to the possible presence of slits or pinholes in the packaging.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=69094 | Z-1481-2008 | Zimmer H.A.S. Autotransfusion System, H.A.S. Kit with 3/16" Soft Drain , 4,8mm, 15Fr, PVC, 107cm, firm, sterile, latex free, Rx only; Zimmer, Swindon, Wiltshire, United Kingdom; Zimmer, Dover, Ohio; REF 00-2555-030-05. | Zimmer H.A.S. Autotransfusion System H.A.S. Kit with 3/16" Drain | 2 | 09/19/2008 | 09/19/2008 | 11/12/2009 | 1000220733 | Zimmer Inc. | The sterility of the device may be compromised due to the possible presence of slits or pinholes in the packaging.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=69097 | Z-1484-2008 | Zimmer H.A.S. Autotransfusion System, H.A.S. Kit without Drain, sterile, latex free, Rx only; Zimmer, Swindon, Wiltshire, United Kingdom; Zimmer, Dover, Ohio; REF 00-2555-060-05. | Zimmer H.A.S. Autotransfusion System Kit w/o Drain | 2 | 09/19/2008 | 09/19/2008 | 11/12/2009 | 1000220733 | Zimmer Inc. | The sterility of the device may be compromised due to the possible presence of slits or pinholes in the packaging.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=69095 | Z-1482-2008 | Zimmer H.A.S. Autotransfusion System, H.A.S. Kit with 1/4" Drain, 6,4mm, 19Fr, PVC, 107cm, large, sterile, latex free, Rx only; Zimmer, Swindon, Wiltshire, United Kingdom; Zimmer, Dover, Ohio; REF 00-2555-040-05. | Zimmer H.A.S. Autotransfusion System H.A.S. Kit with 1/4" Drain | 2 | 09/19/2008 | 09/19/2008 | 11/12/2009 | 1000220733 | Zimmer Inc. | The sterility of the device may be compromised due to the possible presence of slits or pinholes in the packaging.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=48115 | Z-1489-06 | dideco Preassembled Surgical Wash Set, Electa Essential , Cobe part numbers WS55E, WS125E, WS175E, WS225E, Cobe Cardiovascular, Inc. | dideco Surgical blood wash set | 2 | 09/08/2006 | 09/09/2006 | 12/28/2011 | 1718850 | Cobe Cardiovascular, Inc | Foreign particles-manufacturing error may lead to abrasion of bowl seal in blood recovery sets, possibly causing particles in the blood.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=48127 | Z-1491-06 | Cobe STAT PAC Autotransfusion Set (made with Electa Wash Set), Cobe Product Codes: ASCBFE22, ASCBE22, AS9E55, AS9E22, AS9E12, AS4E55, AS4E22, AS4E17. Cobe Cardiovascular Inc. | COBE Autotransfusion set. | 2 | 09/08/2006 | 09/09/2006 | 12/28/2011 | 1718850 | Cobe Cardiovascular, Inc | Foreign particles-manufacturing error may lead to abrasion of bowl seal in blood recovery sets, possibly causing particles in the blood.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=48113 | Z-1488-06 | dideco Preassembled Surgical Wash Set, Compact , Cobe part numbers WS55C, WS125C, WS225C, Cobe Cardiovascular, Inc. | dideco Surgical blood wash set | 2 | 09/08/2006 | 09/09/2006 | 12/28/2011 | 1718850 | Cobe Cardiovascular, Inc | Foreign particles-manufacturing error may lead to abrasion of bowl seal in blood recovery sets, possibly causing particles in the blood.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=48125 | Z-1490-06 | Cobe STAT PAC, Autotransfusion Set (Made with Compact Wash Sets), Cobe Product Codes: AS4C12, AS4C22, ASCBC12C, ASCBC22C, ASCBFC22C. Cobe Cardiovascular Inc. | COBE Autotransfusion set. | 2 | 09/08/2006 | 09/09/2006 | 12/28/2011 | 1718850 | Cobe Cardiovascular, Inc | Foreign particles-manufacturing error may lead to abrasion of bowl seal in blood recovery sets, possibly causing particles in the blood.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=166278 | Z-2950-2018 | Haemonetics OrthoPAT System-Integrated Set. Product Code: 1150H-00 OrthoPAT orthopedic perioperative autotransfusion system indicated for use to salvage RBCs from blood lost intraoperatively and postoperatively during surgical procedures | Haemonetics OrthoPat | 2 | 08/30/2018 |  | 06/16/2020 | 1219343 | Haemonetics Corporation | OrthoPAT reservoirs are missing the cannula inside the reservoirs. The absence of this cannula prevents the emptying of blood from the reservoirs. The inability to empty the reservoir means that the collected blood cannot be processed.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=166279 | Z-2951-2018 | Haemonetics OrthoPAT System-OPAT QuickConnect Reservoirs. Product Code: OPT-R-1000 The OrthoPAT orthopedic perioperative autotransfusion system indicated for use to salvage RBCs from blood lost intraoperatively and postoperatively during surgical procedures | Haemonetics OrthoPat | 2 | 08/30/2018 |  | 06/16/2020 | 1219343 | Haemonetics Corporation | OrthoPAT reservoirs are missing the cannula inside the reservoirs. The absence of this cannula prevents the emptying of blood from the reservoirs. The inability to empty the reservoir means that the collected blood cannot be processed.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=102445 | Z-2956-2011 | Haemonetics cardioPAT -Cardiovascular Perioperative Autotransfusion System Model 2050 US Models Out-Side the US: 02050-UK, 02050-ES, 02050-IT, 02050-NL, 02050-GE, 02050-FR, 02050-JA The Haemonetics Cardiovascular Perioperative Autotransfusion System for Use (cardioPAT) is indicated for use to salvage red blood cells from blood lost intraoperatively and postoperatively during cardiovascular surgical procedures, where the expected rate of processing of salvaged blood and fluid aspirated from the surgical site is less than or equal to two liters per hour. | Haemonetics cardioPAT -Cardiovascular Perioperative Autotransfusion System | 2 | 08/05/2011 | 08/05/2011 | 06/09/2014 | 1219343 | Haemonetics Corporation | Updated Operations Manual for the deployment of the spill collection system and cleaning
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=40344 | Z-1071-05 | ASY Brat 2 Procedure Set | Autotransfusion apparatus | 2 | 08/02/2005 | 08/03/2005 | 07/27/2006 | 1718850 | Cobe Cardiovascular, Inc | Packaging trays may have holes in them as a result of the sealing process, which could compromise the sterility of the product.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=40345 | Z-1072-05 | ASY, Brat Pack Procedure Set | Autotransfusion apparatus | 2 | 08/02/2005 | 08/03/2005 | 07/27/2006 | 1718850 | Cobe Cardiovascular, Inc | Packaging trays may have holes in them as a result of the sealing process, which could compromise the sterility of the product.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=40347 | Z-1074-05 | ASY, Brat 2 Procedure Set | Autotransfusion apparatus | 2 | 08/02/2005 | 08/03/2005 | 07/27/2006 | 1718850 | Cobe Cardiovascular, Inc | Packaging trays may have holes in them as a result of the sealing process, which could compromise the sterility of the product.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=40346 | Z-1073-05 | ASY Brat 2 Procedure Set | Autotransfusion apparatus | 2 | 08/02/2005 | 08/03/2005 | 07/27/2006 | 1718850 | Cobe Cardiovascular, Inc | Packaging trays may have holes in them as a result of the sealing process, which could compromise the sterility of the product.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=40343 | Z-1070-05 | ASY, Brat 2 Pack Procedure Set | Autotransfusion apparatus | 2 | 08/02/2005 | 08/03/2005 | 07/27/2006 | 1718850 | Cobe Cardiovascular, Inc | Packaging trays may have holes in them as a result of the sealing process, which could compromise the sterility of the product.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=40334 | Z-1069-05 | ASY Brat 2 Procedure Set, for blood recovery & autotransfusion. | Autotransfusion apparatus | 2 | 08/02/2005 | 08/03/2005 | 07/27/2006 | 1718850 | Cobe Cardiovascular, Inc | Packaging trays may have holes in them as a result of the sealing process, which could compromise the sterility of the product.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=127513 | Z-1959-2014 | Atrium Express Dry Seal Chest Drain ATS Blood Recovery, Catalog Number: 4050-100N. Autotransfusion Apparatus. | Atrium Express Dry Seal Chest Drain ATS Blood Recovery | 2 | 07/01/2014 | 07/01/2014 | 03/17/2016 | 3011175548 | Atrium Medical Corporation | Product that was previously recalled because the Chest Drain tubing of the ATS Blood recovery may leak or disconnect, was released for distribution in error.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=99797 | Z-2420-2011 | Haemonetics OrthoPAT -Orthopedic Perioperative Autotransfusion System US Only: 1050-110-ENG Outside US: 1050-110-ENG, I050-240-XXX (generic), 1050-240-ENG, 1050-240-SPN, 1050-240-ITL, 1050-240GER, 1050-240-DUT, 1050-240-FRN, 1050-240-CZE, 1050-IIO-JPN. Provides perioperative autotransfusion ( salvaging and reinfusing red blood cells (RBC's) from blood shed from patients undergoing orthopedic surgery. The portable system is used in operating rooms to recycle blood lost during surgery and in the recovery room to recycle blood lost after surgery. It can be transported to the patient floor for postoperative salvage. | ORTHOPEDIC PERIOPERATIVE AUTOTRANSFUSION | 2 | 06/07/2011 | 06/07/2011 | 04/21/2014 | 1219343 | Haemonetics Corporation | 2 Models Model 1: OrthoPAT devices with no centrifuge drain spill containment system may have potential for fluid contamination leading to potential spark or smoke may occur are recalled Model 2: OrthoPat equipped with a spill collection drainage system updated with revised Instructions for Use and Quick reference Guide Addendum
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=25980 | Z-0906-03 | ATF 40 Fast Start Kit - contains one each AT1 Autotransfusion Set, ATS Suction Line, and the ATR40 Autotransfusion Reservoir. The component that is the subject of the recall is the AT1 Autotransfusion Set. | Fresenius | 2 | 06/04/2003 | 06/05/2003 | 06/04/2003 | 3002987955 | Fresenius Hemocare, Inc. | Sterility may be compromised
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135099 | Z-1570-2015 | 3/16 in. Curved Trocar with Wound Drain Curved Trocar with Wound Drain/Double Trocar with Wound Drain: The curved trocar wound drain has a curved metal trocar attached to one end of the flexible wound drain tube. There is also a polymeric trocar guard that provides protection from the sharp end of the trocar during handling. The double trocar wound drain has metal trocars on either side of the flexible wound drain tube. These are also equipped with trocar guards. | CBC II Wound Drains | 2 | 05/01/2015 | 05/01/2015 | 12/31/2015 | 1811755 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the CBCII Wound Drains and Ambulating Bulb for a potential sterility breach in the packaging
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135101 | Z-1572-2015 | 1/4 in. Curved Trocar with Wound Drain Curved Trocar with Wound Drain/Double Trocar with Wound Drain: The curved trocar wound drain has a curved metal trocar attached to one end of the flexible wound drain tube. There is also a polymeric trocar guard that provides protection from the sharp end of the trocar during handling. The double trocar wound drain has metal trocars on either side of the flexible wound drain tube. These are also equipped with trocar guards. | CBC II Wound Drains | 2 | 05/01/2015 | 05/01/2015 | 12/31/2015 | 1811755 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the CBCII Wound Drains and Ambulating Bulb for a potential sterility breach in the packaging
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135100 | Z-1571-2015 | 1/8 in. Curved Trocar with Wound Drain Curved Trocar with Wound Drain/Double Trocar with Wound Drain: The curved trocar wound drain has a curved metal trocar attached to one end of the flexible wound drain tube. There is also a polymeric trocar guard that provides protection from the sharp end of the trocar during handling. The double trocar wound drain has metal trocars on either side of the flexible wound drain tube. These are also equipped with trocar guards. | CBC II Wound Drains | 2 | 05/01/2015 | 05/01/2015 | 12/31/2015 | 1811755 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the CBCII Wound Drains and Ambulating Bulb for a potential sterility breach in the packaging
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=135103 | Z-1573-2015 | Ambulating Bulb Ambulating Bulb: The Ambulating Bulb comprises four main components: the suction reservoir, the PVC tubing, the slide clamp and the female connector. The rigid polycarbonate female connector is used to connect the ambulating bulb to the CBCII evacuator tube in order to collect wound drainage. | CBC II Wound Drains | 2 | 05/01/2015 | 05/01/2015 | 12/31/2015 | 1811755 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the CBCII Wound Drains and Ambulating Bulb for a potential sterility breach in the packaging
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=43656 | Z-0822-06 | Haemonetics Cell Saver 5 Autogolous Blood Recovery System Model Numbr: LN02005-110-E |  | 2 | 04/28/2006 | 04/29/2006 | 03/12/2008 | 1219343 | Haemonetics Corporation | Device emits radiofrequency which exceed the IEC 60601-1-2 International Standard
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=43667 | Z-0823-06 | Haemonetics Cell Saver 5+ Autogolous Blood Recovery System Model Number: LN02005-110EP |  | 2 | 04/28/2006 | 04/29/2006 | 03/12/2008 | 1219343 | Haemonetics Corporation | Device emits radiofrequency which exceed the IEC 60601-1-2 International Standard
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=144723 | Z-1582-2016 | Pleur-evac Chest Drainage System, Dry Suction/Dry Seal Control, REF A-6000-08LF, Rx Only, Teleflex Medical. Product Usage: Pleur-Evacs are sterile, single-use, chest drainage systems intended for post-operative use | Pleur-evac Chest Drainage System | 2 | 04/27/2016 |  | 01/30/2025 | 3005747797 | Teleflex Medical | The label on the Tyvek bag is missing. This is used by customers to identify material code, lot number, and expiration date once the product is removed from the shipper box.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171537 | Z-1209-2019 | Fresenius Kabi CATSmart device (Continuous autotransfusion System). Labeled as the following kits: 1. AT-1 sets and kits (9005104 AT-1 Autotransfusion set, 9108504 ATF 120 Fast Start Kit, and 9108494 ATF 40 Fast Start Kits); 2. AT-3 Autotransfusion sets and kits (Part 9005444 AT-3 set, 9104444 ATF-120 Fast Start Kits (AT-3), or 9108444 ATF-40 Fast Start Kits AT-3) Product Usage: The Fresenius Kabi CATSmart device (Continuous auto transfusion System) is indicated for the processing of autologous shed blood collected interaoperatively and postoperatively to obtain washed packed red blood cells for reinfusion. The CATSmart Auto transfusion System is designed to reduce the need for allogeneic blood transfusion by re-infusing a patient s own lost blood during surgery. The set is used as a disposable component of the autotransfusion system during surgery and/or treatment of trauma to remove waste and improve the safety of autologous re-infused blood (auto transfusion). This is a single-use device. | CATSmart Autotransfusion | 2 | 04/23/2019 |  | 10/05/2020 | 3004548776 | Fenwal Inc | Based on internal investigations, Fresenius Kabi has identified a multifaceted set of events involving the CATSmart Systems, including the device, software, and disposable which may result in the centrifuge tubing line making contact with the washing chamber. The contact may result in excessive wear of the centrifuge tubing or in extreme cases a puncture of the tubing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171116 | Z-1059-2019 | Vacuum Unit for CATSmart Vacuum Pump. Fresenius CATSmart (Continuous Autotransfusion System) is indicated for the processing of blood collected intraoperatively and postoperatively to obtain washed packed red blood cells for reinfusion. | CATSmart Vacuum Pump | 2 | 03/26/2019 |  | 06/23/2020 | 3004548776 | Fenwal Inc | The optional Vacuum Unit may stop working and display the failure message "Failure vacuum unit" during use.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=97477 | Z-1391-2011 | Haemonetics cardioPAT Cardiovascular - autotransfusion System Catalog Number: 02050-US | Haemonetics cardioPAT | 2 | 02/22/2011 | 02/22/2011 | 01/14/2013 | 1219343 | Haemonetics Corporation | Update Instructions for Use for the CardioPAT&reg; Manual for Leak Detection. User may mis-interpret the air-leak rate displayed by the cardioPAT device when making a determination about chest-drain removal, which could result in the patient experiencing a pneumothorax.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=151090 | Z-1207-2017 | Maquet Oasis Dry Suction Water Seal Chest Drain, Single Collection w/AC Reference: 3600-100 Intended Use: To evacuate air and/or fluid from the chest cavity or mediastinum. " To help re-establish lung expansion and restore breathing dynamics. " To facilitate postoperative collection and reinfusion of autologous blood from the patients pleural cavity or mediastinal area | Atrium Medical | 2 | 02/15/2017 |  | 11/03/2020 | 3011175548 | Atrium Medical Corporation | Chest drains manual vent valve located at the top of affected chest drains could potentially not re-seal after being depressed.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=152134 | Z-1211-2017 | Maquet Express Dry Seal Suction OASIS DRAIN, DRY PEDI W/AC Reference: 3612-100 Intended Use: To evacuate air and/or fluid from the chest cavity or mediastinum. " To help re-establish lung expansion and restore breathing dynamics. " To facilitate postoperative collection and reinfusion of autologous blood from the patients pleural cavity or mediastinal area | Atrium Medical | 2 | 02/15/2017 |  | 11/03/2020 | 3011175548 | Atrium Medical Corporation | Chest drains manual vent valve located at the top of affected chest drains could potentially not re-seal after being depressed.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=151091 | Z-1208-2017 | Maquet Express Dry Seal Suction Chest Drain, Single Collection w/AC Reference: 4000-100N Intended Use: To evacuate air and/or fluid from the chest cavity or mediastinum. " To help re-establish lung expansion and restore breathing dynamics. " To facilitate postoperative collection and reinfusion of autologous blood from the patients pleural cavity or mediastinal area | Atrium Medical | 2 | 02/15/2017 |  | 11/03/2020 | 3011175548 | Atrium Medical Corporation | Chest drains manual vent valve located at the top of affected chest drains could potentially not re-seal after being depressed.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=49575 | Z-0500-2007 | Atrium Express Dry Seal Chest Drain with Sterile Fluid Path Package Model Number: 4050-170N | Atrium | 2 | 02/14/2007 | 02/15/2007 | 11/27/2007 | 3011175548 | Atrium Medical Corporation | Chest drainage tubing incorrectly assembled and may disconnect from patient during use
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178149 | Z-0900-2020 | JAPAN DISP SET, 261J, 205J, HAR - Product Usage: its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma. The recovered blood is processed by a centrifugation and washing procedure before pumping the red cell product to either a bag for gravity reinfusion into the patient or an extracorporeal circuit. | CS Bowl Set | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178148 | Z-0899-2020 | JAPAN DISP SET, 261J, 205J, 208 - Product Usage: its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma. The recovered blood is processed by a centrifugation and washing procedure before pumping the red cell product to either a bag for gravity reinfusion into the patient or an extracorporeal circuit. | CS Bowl Set | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178154 | Z-0905-2020 | Japan DISP SET, 263J, 205J, 208 - Product Usage: its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma | CS5 Bowl Kit | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178147 | Z-0898-2020 | JAPAN BASIC HI SPEED CS5 SET,LN 263J - Product Usage: its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma. The recovered blood is processed by a centrifugation and washing procedure before pumping the red cell product to either a bag for gravity reinfusion into the patient or an extracorporeal circuit. | CS Bowl Set | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178151 | Z-0902-2020 | JAPAN CS ELITE FASTPACK, 225ML - Product Usage: its related accessory components are intended for use to recover blood shed during or subsequent to an operation or as a result of trauma, processing the blood by a centrifugation and washing procedure, and pumping this processed red cell product to either a bag for gravity reinfusion into the patient or to the arterial line of an extracorporeal circuit for reinfusion into the patient. | CSE Fastpack Set | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178155 | Z-0906-2020 | Japan DISP SET, 263J, 205J, HAR - Product Usage: its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma | CS5 Set | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=177531 | Z-0890-2020 | CS5/5+ Fastpacks, 225 mL, 20 Res Autotransfusion Device - Product Usage: its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma. | CS5/5+ Fastpacks | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=177534 | Z-0893-2020 | CS Elite Processing Set (125 ml) Autotransfusion Device - Product Usage:its related accessory components are intended for use to recover blood shed during or subsequent to an operation or as a result of trauma, processing the blood by a centrifugation and washing procedure, and pumping this processed red cell product to either a bag for gravity reinfusion into the patient or to the arterial line of an extracorporeal circuit for reinfusion into the patient. | CS Elite | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178150 | Z-0901-2020 | Japan CS Elite FASTPACK DISP. SET - 125ML - Product Usage: its related accessory components are intended for use to recover blood shed during or subsequent to an operation or as a result of trauma, processing the blood by a centrifugation and washing procedure, and pumping this processed red cell product to either a bag for gravity reinfusion into the patient or to the arterial line of an extracorporeal circuit for reinfusion into the patient. | CSE Fastpack Set | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=177528 | Z-0887-2020 | High Speed Cell Saver5/5+ Bowl Set Autotransfusion Device - Product Usage: its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma. | High Speed Cell Saver | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178152 | Z-0903-2020 | JAPAN CELL SAVER ELITE SET - 125ML - Product Usage: its related accessory components are intended for use to recover blood shed during or subsequent to an operation or as a result of trauma, processing the blood by a centrifugation and washing procedure, and pumping this processed red cell product to either a bag for gravity reinfusion into the patient or to the arterial line of an extracorporeal circuit for reinfusion into the patient. | CSE Set | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=177525 | Z-0885-2020 | CS5/5+ Fastpacks, 225 mL, 150 Res Autotransfusion Device - Product Usage: its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma. | CS5/5+ Fastpacks | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=177530 | Z-0889-2020 | CS5/5+ Fastpacks, 125 mL, 20 Res Autotransfusion Device- Product Usage: its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma. | CS5/5+ Fastpacks | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=177532 | Z-0891-2020 | Cell Saver5/5+ Bowl Kit (225 ml) Autotransfusion Device - Product Usage: its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma. | Cell Saver | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=177536 | Z-0895-2020 | CS Elite Fastpack, 225 mL, 150 Res Autotransfusion Device - Product Usage: its related accessory components are intended for use to recover blood shed during or subsequent to an operation or as a result of trauma, processing the blood by a centrifugation and washing procedure, and pumping this processed red cell product to either a bag for gravity reinfusion into the patient or to the arterial line of an extracorporeal circuit for reinfusion into the patient. | CS Elite Fastpack | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=177537 | Z-0896-2020 | CS Elite Fastpack, 125 mL, 150 Res Autotransfusion Device - Product Usage: its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma. | CS Elite Fastpack | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178146 | Z-0897-2020 | JAPAN LOW VOL(125ML) CS5 SET,LN 261J - Product Usage: its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma. The recovered blood is processed by a centrifugation and washing procedure before pumping the red cell product to either a bag for gravity reinfusion into the patient or an extracorporeal circuit. | CS Set | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=177529 | Z-0888-2020 | CS5/5+ Fastpacks, 125 mL, 150 Res Autotransfusion Device - Product Usage: its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma. | CS5/5+ Fastpacks | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=177535 | Z-0894-2020 | CS Elite Fastpack, 225 mL, 20 Res Autotransfusion Device - Product Usage: its related accessory components are intended for use to recover blood shed during or subsequent to an operation or as a result of trauma, processing the blood by a centrifugation and washing procedure, and pumping this processed red cell product to either a bag for gravity reinfusion into the patient or to the arterial line of an extracorporeal circuit for reinfusion into the patient. | CS Elite Fastpack | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178153 | Z-0904-2020 | JAPAN CELL SAVER ELITE SET - 225ML - Product Usage: its related accessory components are intended for use to recover blood shed during or subsequent to an operation or as a result of trauma, processing the blood by a centrifugation and washing procedure, and pumping this processed red cell product to either a bag for gravity reinfusion into the patient or to the arterial line of an extracorporeal circuit for reinfusion into the patient. | CSE Set | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=177527 | Z-0886-2020 | CS5/5+ Bowl Kit-125 mL Autotransfusion Device - Product Usage: its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma. | CS5/5+ Bowl Kit | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=177533 | Z-0892-2020 | CS Elite Processing Set (225 ml) Autotransfusion Device - Product Usage:its related accessory components are intended for use to recover blood shed during or subsequent to an operation or as a result of trauma, processing the blood by a centrifugation and washing procedure, and pumping this processed red cell product to either a bag for gravity reinfusion into the patient or to the arterial line of an extracorporeal circuit for reinfusion into the patient. | CS Elite | 2 | 01/29/2020 |  | 11/06/2020 | 1219343 | Haemonetics Corporation | Potential to develop leaks at the inner core of Cell Saver 5/5+ and Cell Saver Elite/Elite +125ml and 225ml bowl sets. Which could result in fluid being trapped inside the bowl core, causing a long empty device error message with the possibility of incomplete washing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=169578 | Z-0785-2019 | RESERVOIR "Y" ADAPTER, (a) BT920 (b) BT926 Suction and Anticoagulation assemblies are indicated for use with an autotransfusion system between the suction tip and reservoir. | RESERVOIR "Y" ADAPTER | 2 | 01/29/2019 |  |  | 1000116158 | Medtronic Perfusion Systems | Medtronic discovered that certain Y-connectors included in the suction lines have exhibited tearing and cracking. These defects could cause the Y-connectors to leak when in use, which could result in potential blood loss and a breach in the circuit.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=169580 | Z-0786-2019 | SUCTION / ANTICOAGULATION ASSEMBLY (a) BTC93 (b) BTC98 Suction and Anticoagulation assemblies are indicated for use with an autotransfusion system between the suction tip and reservoir. | SUCTION / ANTICOAGULATION ASSEMBLY | 2 | 01/29/2019 |  |  | 1000116158 | Medtronic Perfusion Systems | Medtronic discovered that certain Y-connectors included in the suction lines have exhibited tearing and cracking. These defects could cause the Y-connectors to leak when in use, which could result in potential blood loss and a breach in the circuit.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87671 | Z-0742-2010 | Stryker CBCII Collection only pkg/6 with 1/8 in Trocar, sterile, Stryker Instruments, Kalamazoo, MI, REF 0225-028-S21. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII Collection only pkg/6 with 1/8 in Trocar | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87692 | Z-0743-2010 | Stryker CBC ConstaVac Blood Conservation Kit 3/16in. (4.8mm) PVC, sterile, Stryker Instruments, Kalamazoo, MI REF 225-426-W. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBC ConstaVac Blood Conservation Kit 3/16in. (4.8mm) PVC, sterile | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87638 | Z-0723-2010 | Stryker CBCII with Quick Disconnect, 3/16 in. Double Trocar and Reinfusion Labels, sterile Stryker Instruments, Kalamazoo, MI, REF 225-28-85S4. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII with Quick Disconnect, 3/16 in. Double Trocar and Reinfusion Labels | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87640 | Z-0725-2010 | Stryker CBCII Blood Conservation Kit w/ 1/8 in. (3.2mm) Round PVC Drain and Curved Trocar, sterile, Stryker Instruments, Kalamzoo, MI; REF 225-28-514. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII Blood Conservation Kit w/ 1/8 in. (3.2mm) Round PVC Drain and Curved Trocar, | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87641 | Z-0726-2010 | Stryker CBCII Blood Conservation Kit w/ 1/4 in. (6.4mm) Round PVC Drain and Curved Trocar, sterile, Stryker Instruments, Kalamzoo, MI, REF 225-28-516. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII Blood Conservation Kit w/ 1/4 in. (6.4mm) Round PVC Drain and Curved Trocar | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87650 | Z-0735-2010 | Stryker CBCII ConstaVac Blood Collection Only System with 1/8in. (3.2mm) PVC Drains and Double 1/8in. (3.2mm) Trocar, sterile, Stryker Instruments, Kalamazoo, MI REF 225-29. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII ConstaVac Blood Collection Only System with 1/8in. (3.2mm) PVC Drains and Double 1/8in | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87654 | Z-0738-2010 | Stryker CBCII ConstaVac Blood Conservaion Kit 1/8in. (3.2mm) PVC, sterile, Stryker Instruments, Kalamazoo, MI, REF 225-414-W. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII ConstaVac Blood Conservaion Kit 1/8in. (3.2mm) PVC | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87643 | Z-0728-2010 | Stryker CBCII Double Trocar Kit with 1/8 in. (3.2mm) Drain, sterile, Stryker Instruments, Kalamazoo, MI, REF 225-28-614. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII Double Trocar Kit with 1/8 in. (3.2mm) Drain, | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87637 | Z-0722-2010 | Stryker CBCII 1/8in. Double Trocar with quick disconnect , sterile, Stryker Instuments, Kalamazoo, MI, REF225-28-85S1. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII 1/8in. Double Trocar with quick disconnect | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87642 | Z-0727-2010 | Stryker CBCII Blood Conservation Kit w 3/16in. (4.8mm) round PVC Drain and Curved Trocar, sterile, Stryker Instruments, Kalamazoo, MI, REF 225-28-526. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII Blood Conservation Kit w/ 13/16 (4.8mm Round PVC Drain and Curved Trocar | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87646 | Z-0731-2010 | Stryker CBCII Double Trocar/Double Pouch Kit with 3/16in. (4.8mm) Drain, sterile, Stryker Instruments, Kalamazoo, MI REF 225-28-626S1. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII Double Trocar/Double Pouch Kit with 3/16in. (4.8mm) Drain | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87658 | Z-0740-2010 | Stryker CBCII ConstaVac Blood Conservation Kit 3/16in. (4.8 mm) PVC, sterile, Stryker Instruments, Kalamazoo, MI, REF 225-426. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII ConstaVac Blood Conservation Kit 3/16in. (4.8 mm) PVC | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87636 | Z-0721-2010 | Stryker CBC Blood Conservation Kit with Quick Disconnect Blood Bag, sterile, Stryker Instruments, Kalamazoo, MI, REF 225-28-85. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBC Blood Conservation Kit with Quick Disconnect Blood Bag | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87639 | Z-0724-2010 | Stryker CBCII with Quick Disconnect and 3/16 in. Double Trocar, sterile, Stryker Instruments, Kalamazoo, MI, REF 225-28-S5. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII with Quick Disconnect and 3/16 in. Double Trocar | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87644 | Z-0729-2010 | Stryker CBCII Double Trocar Kit with 1/4in. (6.4mm) Drain, sterile, Stryker Instruments, Kalamazoo, MI, REF 225-28-616. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII Double Trocar Kit with 1/4in. (6.4mm) Drain | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87645 | Z-0730-2010 | Stryker CBCII Double Trocar Kit with 3/16in. (4.8mm) Drain, sterile, Stryker Instruments, Kalamazoo, MI, REF 225-28-626. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII Double Trocar Kit with 3/16in. (4.8mm) Drain | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87648 | Z-0733-2010 | Stryker CBCII ConstaVac Blood Conservation System with Quick Disconnect and 1/4in. (6.4mm) PVC Drain, sterile, Stryker Instruments, Kalamazoo, MI REF 225-28-916. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII ConstaVac Blood Conservation System with Quick Disconnect and 1/4in. (6.4mm) PVC Drain | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87632 | Z-0719-2010 | Stryker CBCII ConstaVac Blood conservation system, sterile, Stryker Instruments, Kalamazoo, MI, REF 225-28E. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII ConstaVac Blood conservation system | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87634 | Z-0720-2010 | Stryker CBCII collection only pkg/6, sterile, Stryker Instruments, Kalamazoo, MI, REF 225-28-S20. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII collection only pkg/6 | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87647 | Z-0732-2010 | Stryker CBCII ConstaVac Blood Conservagtion System w/ Quick Disconnect and 1/8in. (3.2mm) PVC Drain, sterile, Stryker Instruments, Kalamazoo, MI REF 225-28-914. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII ConstaVac Blood Conservagtion System w/ Quick Disconnect and 1/8in. (3.2mm) PVC Drain | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87652 | Z-0737-2010 | Stryker CBCII ConstaVac Blood Conservation Kit 1/8in. (3.2mm) PVC w/o Fat Retention Valve, sterile, Stryker Instruments, Kalamazoo, MI, REF 225-414S2. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII ConstaVac Blood Conservation Kit 1/8in. (3.2mm) PVC w/o Fat Retention Valve | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87665 | Z-0741-2010 | Stryker CBCII with 3/16 in. Silicone Drain, sterile, Stryker Instruments, Kalamazoo, MI, REF 0225-028-S16. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII with 3/16 in. Silicone Drain | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87651 | Z-0736-2010 | Stryker CBCII ConstaVAc Blood Conservatin Kit 1/8in. (3.2mm) PVC, sterile, Stryker Instruments, Kalamazoo, MI, REF 225-414. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII ConstaVAc Blood Conservatin Kit 1/8in. (3.2mm) PVC, | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87656 | Z-0739-2010 | Stryker CBCII ConstaVac Blood Conservation Kit 1/4in. (6.4mm) PVC, sterile, Stryker Instruments, Kalamazoo, MI, REF 225-416. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII ConstaVac Blood Conservation Kit, | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=87649 | Z-0734-2010 | Stryker CBCII ConstaVac Blood Conservation System w/Quick Disconnect and 3/16in. (4.8mm) PVC Drain, sterile, Stryker Instruments, Kalamazoo, MI, REF 225-28-926. Blood recovery system used post-operatively to collect, filter and allow for reinfusion of autologous blood. | Stryker CBCII ConstaVac Blood Conservation System w/Quick Disconnect and 3/16in. (4.8mm) PVC Drain | 2 | 01/29/2010 | 01/29/2010 | 10/25/2010 | 1811755 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=25630 | Z-0501-03 | Haemonetics High Speed Cell Saver Bowl Set List Number: 263 | Disposable Set for autologous blood cell recovery | 3 | 01/29/2003 | 01/30/2003 | 01/28/2016 | 1219343 | Haemonetics Corporation | Disposable Set for Autologous Blood recovery may leak saline at the Y-connector
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=25629 | Z-0500-03 | Haemonetics Basic Low Volume (125ml) Cell Saver 5 Bowl Set List Number: 261 | Disposable Set for autologous blood cell recovery | 3 | 01/29/2003 | 01/30/2003 | 01/28/2016 | 1219343 | Haemonetics Corporation | Disposable Set for Autologous Blood recovery may leak saline at the Y-connector
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=25632 | Z-0503-03 | Haemonetics Cell Saver 5 Low Volume (12ml) Fast Pack with Filtered Reservoir List Number: 265F | Disposable Set for autologous blood cell recovery | 3 | 01/29/2003 | 01/30/2003 | 01/28/2016 | 1219343 | Haemonetics Corporation | Disposable Set for Autologous Blood recovery may leak saline at the Y-connector
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=25626 | Z-0498-03 | Haemonetics Cell Saver 5 Fast Pack List Number: 260 | Disposable Set for autologous blood cell recovery | 3 | 01/29/2003 | 01/30/2003 | 01/28/2016 | 1219343 | Haemonetics Corporation | Disposable Set for Autologous Blood recovery may leak saline at the Y-connector
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=25628 | Z-0499-03 | Haemonetics Cell Saver 5 Fast Pack with Filtered Reservoir List Number: 260F | Disposable Set for autologous blood cell recovery | 3 | 01/29/2003 | 01/30/2003 | 01/28/2016 | 1219343 | Haemonetics Corporation | Disposable Set for Autologous Blood recovery may leak saline at the Y-connector
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=25631 | Z-0502-03 | Haemonetics Cell Saver 5 Low Volume (125mL) Fast Pack List Number: 265 | Disposable Set for autologous blood cell recovery | 3 | 01/29/2003 | 01/30/2003 | 01/28/2016 | 1219343 | Haemonetics Corporation | Disposable Set for Autologous Blood recovery may leak saline at the Y-connector
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161214 | Z-0779-2019 | Cell Saver Elite Processing Kit, 70 ml, Product ID: CSE-P-70 Indications for Use: The Haemonetics Cell Saver 5/5+ System and its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma. | Autotransfusion Apparatus | 2 | 01/28/2019 |  | 08/06/2020 | 1219343 | Haemonetics Corporation | Haemonetics has received isolated reports that in some cases the 70ml centrifuge bowl may leak during cell salvage processing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161213 | Z-0778-2019 | Cell Saver 5/5+ Bowl Set, 70ml, Product ID: 0291E-00, 0291A-00 Indications for Use: The Haemonetics Cell Saver 5/5+ System and its related accessory components are intended to be used for the recovery and processing of blood shed during or subsequent to an operation or as a result of trauma. | Autotransfusion Apparatus | 2 | 01/28/2019 |  | 08/06/2020 | 1219343 | Haemonetics Corporation | Haemonetics has received isolated reports that in some cases the 70ml centrifuge bowl may leak during cell salvage processing.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=123497 | Z-0741-2014 | Atrium Oasis Dry Suction Water Seal ATS (Autotransfusion) Blood Recovery System Chest Drain Dual patient lines; Product Code: 3652-100. To evacuate air and/or fluid from the chest cavity or mediastinum. | Atrium Medical Oasis Dry Suction Water Seal ATS | 2 | 01/14/2014 | 01/14/2014 | 06/11/2015 | 3011175548 | Atrium Medical Corporation | Chest Drain tubing of of the ATS Blood recovery may leak or disconnect
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=123427 | Z-0735-2014 | Atrium Ocean Water Seal Chest Drain ATS (Autotransfusion) Blood Recovery System with stopcock and in-line connectors; Product Code: 2050-000. To evacuate air and/or fluid from the chest cavity or mediastinum. | Atrium Medical Ocean Water Seal Chest Drain ATS | 2 | 01/14/2014 | 01/14/2014 | 06/11/2015 | 3011175548 | Atrium Medical Corporation | Chest Drain tubing of of the ATS Blood recovery may leak or disconnect
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=123495 | Z-0739-2014 | Atrium Ocean Water Seal Chest Drain ATS (Autotransfusion) Blood Recovery System Dual patient lines with stopcock; .Product Code: 2052-300. To evacuate air and/or fluid from the chest cavity or mediastinum. | Atrium Medical Ocean Water Seal Chest Drain ATS | 2 | 01/14/2014 | 01/14/2014 | 06/11/2015 | 3011175548 | Atrium Medical Corporation | Chest Drain tubing of of the ATS Blood recovery may leak or disconnect
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=123498 | Z-0742-2014 | Atrium Express Dry Seal ATS (Autotransfusion) Blood Recovery System Chest Drain; Product Code: 4050-100N. To evacuate air and/or fluid from the chest cavity or mediastinum. | Atrium Medical Express Dry Seal ATS | 2 | 01/14/2014 | 01/14/2014 | 06/11/2015 | 3011175548 | Atrium Medical Corporation | Chest Drain tubing of of the ATS Blood recovery may leak or disconnect
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=123499 | Z-0743-2014 | Atrium Express Dry Seal ATS (Autotransfusion) Blood Recovery System Chest Drain with Prefilled Air Leak Monitor; Product Code: 4050-100P. To evacuate air and/or fluid from the chest cavity or mediastinum. | Atrium Medical Express Dry Seal ATS | 2 | 01/14/2014 | 01/14/2014 | 06/11/2015 | 3011175548 | Atrium Medical Corporation | Chest Drain tubing of of the ATS Blood recovery may leak or disconnect
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=123493 | Z-0737-2014 | Atrium Ocean Water Seal Chest Drain ATS (Autotransfusion) Blood Recovery System with stopcock Blood Recovery System; Product Code: 2050-300. To evacuate air and/or fluid from the chest cavity or mediastinum. | Atrium Medical Ocean Water Seal Chest Drain ATS | 2 | 01/14/2014 | 01/14/2014 | 06/11/2015 | 3011175548 | Atrium Medical Corporation | Chest Drain tubing of of the ATS Blood recovery may leak or disconnect
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=123494 | Z-0738-2014 | Atrium Ocean Water Seal Chest Drain ATS (Autotransfusion) Blood Recovery System Dual patient lines with stopcock and in-line connectors Product Code: 2052-000. To evacuate air and/or fluid from the chest cavity or mediastinum. | Atrium Medical Ocean Water Seal Chest Drain ATS | 2 | 01/14/2014 | 01/14/2014 | 06/11/2015 | 3011175548 | Atrium Medical Corporation | Chest Drain tubing of of the ATS Blood recovery may leak or disconnect
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=123496 | Z-0740-2014 | Atrium Oasis Dry Suction Water Seal ATS (Autotransfusion) Blood Recovery System Chest Drain with in-line connectors Product Code: 3650-100. To evacuate air and/or fluid from the chest cavity or mediastinum. | Atrium Medical Oasis Dry Suction Water Seal ATS | 2 | 01/14/2014 | 01/14/2014 | 06/11/2015 | 3011175548 | Atrium Medical Corporation | Chest Drain tubing of of the ATS Blood recovery may leak or disconnect
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=123492 | Z-0736-2014 | Atrium Ocean Water Seal Chest Drain ATS Blood Recovery System with stopcock and in-line connectors Sterile fluid path packaging Product Code: 2050-070. To evacuate air and/or fluid from the chest cavity or mediastinum. | Atrium Medical Ocean Water Seal Chest Drain ATS | 2 | 01/14/2014 | 01/14/2014 | 06/11/2015 | 3011175548 | Atrium Medical Corporation | Chest Drain tubing of of the ATS Blood recovery may leak or disconnect
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=204705 | Z-0710-2024 | Product REF 4050-100N, DRAIN, EXPRESS BRU W/AC, Blood Recovery Unit | Atrium Express Dry Suction Dry Seal Chest Drain | 1 | 01/12/2024 | 01/12/2024 |  | 2242352 | Maquet Cardiovascular, LLC | Express chest drains are pre-packaged with sterile water syringes which have bene recalled. If the device is not kept in an upright position, the patient may be exposed to infectious material. Subsequent hazards include but are not limited to infection, abscess, sepsis, and death.
2 | CAC | APPARATUS, AUTOTRANSFUSION | 868.5830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=204704 | Z-0709-2024 | Product REF 4000-100N, DRAIN, EXPRESS SINGLE W/AC, Single Collection | Atrium Express Dry Suction Dry Seal Chest Drain | 1 | 01/12/2024 | 01/12/2024 |  | 2242352 | Maquet Cardiovascular, LLC | Express chest drains are pre-packaged with sterile water syringes which have bene recalled. If the device is not kept in an upright position, the patient may be exposed to infectious material. Subsequent hazards include but are not limited to infection, abscess, sepsis, and death.
2
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130110 | Z-0045-2015 | Covidien Kendall Adult Multi-Function Defibrillation Electrodes Part Number: 20660 | Kendall | 1 | 10/16/2014 | 10/15/2014 | 05/24/2017 | 1282497 | Covidien LLC | Covidien Medi-Trace Cadence and Kendall Defibrillation Electrodes Not Compatible for use with Philips model FR3 or FRx AEDs
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130109 | Z-0044-2015 | Coviden Medi-Trace Cadence Adult Multi-Function Defibrillation Electrodes, Pre-connected Part Number: 22660PC | Covidien | 1 | 10/16/2014 | 10/15/2014 | 05/24/2017 | 1282497 | Covidien LLC | Covidien Medi-Trace Cadence and Kendall Defibrillation Electrodes Not Compatible for use with Philips model FR3 or FRx AEDs
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130108 | Z-0043-2015 | Covidien Medi-Trace Cadence and Kendall Adult Multi-Function Defibrillation Electrodes, Radiotransparent Part Number: 22660R | Covidien | 1 | 10/16/2014 | 10/15/2014 | 05/24/2017 | 1282497 | Covidien LLC | Covidien Medi-Trace Cadence and Kendall Defibrillation Electrodes Not Compatible for use with Philips model FR3 or FRx AEDs Defibrillators
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=130148 | Z-0048-2015 | Philips HEARTSTART Multifunction Electrode Pads Part Number: M3718A | Philips | 1 | 10/16/2014 | 10/15/2014 | 05/24/2017 | 1282497 | Covidien LLC | Covidien Medi-Trace Cadence and Kendall Defibrillation Electrodes Not Compatible for use with Philips model FR3 or FRx AEDs
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189217 | Z-0092-2022 | medex LogiCal 5mL Closed Blood Sampling System, Model Numbers: a) MX961E181P1 b) MX961E211P1 c) MX962E211P1 d) MX96AE181P1 e) MX96AE211P1 | medex LogiCal 5mL Closed Blood Sampling System | 2 | 10/13/2021 |  | 09/22/2023 | 3012307300 | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189223 | Z-0097-2022 | smiths medical medex CBSS 5 ml, einfach proximaler Entnahmeport 60 cm, Model Number DPSHC0083 | smiths medical medex CBSS 5 ml, einfach proximaler Entnahmeport | 2 | 10/13/2021 |  | 09/22/2023 | 3012307300 | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189229 | Z-0103-2022 | smiths medical medex CBSS-5ML, Double set, (172+15cm), Model Number DPSHC0094 | smiths medical medex CBSS-5ML, Double set | 2 | 10/13/2021 |  | 09/22/2023 | 3012307300 | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189222 | Z-0096-2022 | smiths medical medex CBSS 5 ml, dreifach proximaler Entnahmeport 60 cm, Model number DPSHC0082 | smiths medical medex CBSS 5 ml, dreifach proximaler Entnahmeport | 2 | 10/13/2021 |  | 09/22/2023 | 3012307300 | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189224 | Z-0098-2022 | smiths medical medex CBSS 5 ml, zweifach proximaler Entnahmeport 60 cm, Model Number DPSHC0084 | smiths medical medex CBSS 5 ml, zweifach proximaler Entnahmeport | 2 | 10/13/2021 |  | 09/22/2023 | 3012307300 | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189221 | Z-0095-2022 | smiths medical medex CBSS 5 ml, Antishunt proximaler Entnahmeport 60 cm, Model Number DPSHC0081 | smiths medical medex CBSS 5 ml, Antishunt proximaler Entnahmeport | 2 | 10/13/2021 |  | 09/22/2023 | 3012307300 | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189226 | Z-0100-2022 | smiths medical medex CBSS-5ML Antishunt set, (172/15cm), Model Number DPSHC0091 | smiths medical medex CBSS-5ML Antishunt set | 2 | 10/13/2021 |  | 09/22/2023 | 3012307300 | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189225 | Z-0099-2022 | smiths medical medex CBSS-5ML Zweifach set, Model Number DPSHC0090 | smiths medical medex CBSS-5ML Zweifach set | 2 | 10/13/2021 |  | 09/22/2023 | 3012307300 | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189227 | Z-0101-2022 | smiths medical medex CBSS-5ML, Triple set, (172/15cm), Model Number DPSHC0092 | smiths medical medex CBSS-5ML, Triple set | 2 | 10/13/2021 |  | 09/22/2023 | 3012307300 | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189228 | Z-0102-2022 | smiths medical medex CBSS-5ML, Single Set, (172/15cm), Model Number DPSHC0093 | smiths medical medex CBSS-5ML, Single Set | 2 | 10/13/2021 |  | 09/22/2023 | 3012307300 | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189219 | Z-0093-2022 | smiths medical medex HemoDraw Plus Closed Blood Sampling Set with LogiCal Pressure Monitoring System, Model Numbers: a) MX961H182P1 b) MX961H212P1 c) MX962H182P1 d) MX962H212P1 | smiths medical medex HemoDraw Plus Closed Blood Sampling Set with LogiCal Pressure Monitoring System | 2 | 10/13/2021 |  | 09/22/2023 | 3012307300 | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=189220 | Z-0094-2022 | smiths medical medex CBSS 2-fach Druckmesset mit Option, Model Number DPSHC0076 | smiths medical medex CBSS 2-fach Druckmesset mit Option | 2 | 10/13/2021 |  | 09/22/2023 | 3012307300 | Smiths Medical ASD Inc. | Reported pressure shifts and incorrect pressure readings in LogiCal Closed Blood Sampling System sets.
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139462 | Z-2546-2015 | MediTrace Cadence Adult Zoll, Preconnect Defibrillation Electrode, Nonsterile Item Code:22770PC | Medtronic | 2 | 09/02/2015 | 09/02/2015 | 05/24/2017 | 1219930 | Covidien LLC | Multi-Function Defibrillation Electrodes, Pre connect- potential for damage to the wire insulation may result in a potentially increased risk for reduced or no patient therapy and patient and/or clinician burns
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=139461 | Z-2545-2015 | MediTrace Cadence Adult Multi-Function Defibrillation Electrodes Item Code:22550PC | Medtronic MediTrace Cadence Adult Zoll, PreconnectDefibrillation Electrode | 2 | 09/02/2015 | 09/02/2015 | 05/24/2017 | 1219930 | Covidien LLC | Multi-Function Defibrillation Electrodes, Pre connect- potential for damage to the wire insulation may result in a potentially increased risk for reduced or no patient therapy and patient and/or clinician burns
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=117920 | Z-1438-2013 | adult defibrillation, cardioversion, pacing, monitoring Disposable Electrodes, Model P-211-M1 and P-214-M1. Product Usage: Used by trained professionals in hospitals, doctors offices, and Emergency Medical Services for adult low-energy defibrillation, transcutaneous pacing, cardioversion and monitoring. | Medtronic/Physio Control Transparent Defibrillation Electrodes | 2 | 05/30/2013 | 05/30/2013 | 09/05/2013 | 1000127185 | TZ Medical Inc. | TZ Medical Inc is recalling its P-211-M1 and P-214-M1 Medtronic / Physio Control Compatible transparent defibrillation electrodes because it may contain a defective white wire on the connector assembly. If a failure were to occur, the end user would be able to identify the failure by inspecting the white wire 1/4 inch from the Gray Medtronic / Physio adapter. The consignees will notice discolor
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=117214 | Z-1184-2013 | Medi-Trace Cadence Adult Radiotransparent Defibrillation Electrode Product ID: 22550R. Intended for use in cardioversion, pacing, ECG monitoring, and defibrillation procedures. | Covidien Medi-Trace Cadence Adult Radiotransparent Defibrillation Electrode | 2 | 04/29/2013 | 04/29/2013 | 11/19/2015 | 1282497 | Covidien LLC | Defibrillation Electrodes may arc/spark result in thermal damage to the leadwire, which could render the device incapable of delivering the appropriate energy or shock to the patient
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=117215 | Z-1185-2013 | Medi-Trace Cadence Pediatric Radiotransparent Defibrillation Electrode Product ID: 22550P. Intended for use in cardioversion, pacing, ECG monitoring, and defibrillation procedures. | Covidien Medi-Trace Cadence Pediatric Radiotransparent Defibrillation Electrode | 2 | 04/29/2013 | 04/29/2013 | 11/19/2015 | 1282497 | Covidien LLC | Defibrillation Electrodes may arc/spark result in thermal damage to the leadwire, which could render the device incapable of delivering the appropriate energy or shock to the patient
2 | DRO | PACEMAKER, CARDIAC, EXTERNAL TRANSCUTANEOUS (NON-INVASIVE) | 870.5550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=94552 | Z-0937-2011 | Welch Allyn PIC 50 Portable Intensive Care System. Model 971081 series. Designed for use in and out of the hospital. | Welch Allyn PIC 50 Portable Intensive Care System | 2 | 01/24/2011 | 01/24/2011 | 07/25/2011 | 3023750 | Welch Allyn Protocol, Inc | Potential likelihood that the Portable Intensive Care (PIC) 50 Defibrillator's screen display may fail. The loss of the ability to view the patient's Electrocardiogram (ECG) on the screen could cause the user to delay or fail to administer defibrillating shocks needed to resuscitate a patient.
2
2 | DSS | CLIP, VASCULAR | 870.3250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180809 | Z-2034-2020 | STEALTH SPRING CLIP 6mm latis (1/2 Force). REF/UDI::A1601/(01)00607915110567 - Product Usage: The device is indicated for surgical clamping during cardiovascular, peripheral vascular, and general surgery. | Vascular Clip | 2 | 05/15/2020 |  | 04/20/2022 | 2027111 | Applied Medical Resources Corp | Failure of cardiovascular clips to provide sufficient occlusion of the vessel.
2 | DSS | CLIP, VASCULAR | 870.3250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180811 | Z-2036-2020 | STEALTH SPRING CLIP 6mm latis (3/4 Force). REF/UDI::A1603/(01)00607915110581 - Product Usage: The device is indicated for surgical clamping during cardiovascular, peripheral vascular, and general surgery. | Vascular Clip | 2 | 05/15/2020 |  | 04/20/2022 | 2027111 | Applied Medical Resources Corp | Failure of cardiovascular clips to provide sufficient occlusion of the vessel.
2 | DSS | CLIP, VASCULAR | 870.3250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180810 | Z-2035-2020 | STEALTH SPRING CLIP 6mm latis (1/4 Force). REF/UDI::A1602/(01)00607915110574 - Product Usage: The device is indicated for surgical clamping during cardiovascular, peripheral vascular, and general surgery. | Vascular Clip | 2 | 05/15/2020 |  | 04/20/2022 | 2027111 | Applied Medical Resources Corp | Failure of cardiovascular clips to provide sufficient occlusion of the vessel.
2 | DSS | CLIP, VASCULAR | 870.3250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180812 | Z-2037-2020 | STEALTH SPRING CLIP 6mm latis (1/4 Force). Non-Sterile, REF/UDI::A1702/(01)10607915110649 - Product Usage: The device is indicated for surgical clamping during cardiovascular, peripheral vascular, and general surgery. | Vascular Clip | 2 | 05/15/2020 |  | 04/20/2022 | 2027111 | Applied Medical Resources Corp | Failure of cardiovascular clips to provide sufficient occlusion of the vessel.
2
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=141378 | Z-0361-2016 | Model 5001102 MPS Delivery Set w/ arrest agent and additive cassettes, heat exchanger &10 ft. delivery tubing; Model 5001102-AS MPS Delivery Set with 6 ft. delivery tubing; Model 7001102 MPS Low Volume Delivery Set for use with cardiopulmonary bypass equipment. | MPS Delivery Set | 2 | 12/02/2015 | 12/02/2015 | 11/17/2016 | 3001665800 | Quest Medical, Inc. | The products have been found to intermittently exhibit a seal failure during use.
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=98620 | Z-0156-2015 | Terumo Custom Cardioplegia Delivery Set - Heat Exchanger Cardiopulmonary Bypass Item Numbers: 5852, 16010 Product Usage: This is a Cardioplegia Delivery System that is designed to delivery blood/solution to the patient undergoing cardiopulmonary bypass surgery. The set is designed to deliver blood/solutions at a 4:1 ratio. In includes a Conducer Heat Exchanger and a Bubble Trap module. Additionally, the set is constructed with a female luer connector that is the subject of the correction/removal activity | Modified MP-4Cardioplegia Delivery System with Conducer | 2 | 10/31/2014 | 10/31/2014 | 11/06/2014 | 1000149028 | Terumo Cardiovascular Systems Corp | Defect in the luer connector can compromise the seal on the female side of the connector and the compromised seal can introduce air in the table line
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=126931 | Z-0161-2015 | Terumo Custom Cardioplegia Delivery Sets Catalog Code: 204880 Product Usage: This is a replacement line to multiple Cardioplegia Delivery Sets-and includes the female luer connector" that is the subject of this correction/removal activity | MiniPrime Cardioplegia Set with Conducer | 2 | 10/31/2014 | 10/31/2014 | 11/06/2014 | 1000149028 | Terumo Cardiovascular Systems Corp | Defect in the luer connector can compromise the seal on the female side of the connector and the compromised seal can introduce air in the table line
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=98879 | Z-0158-2015 | Terumo Cardioplegia Delivery Set- Item Number: 16130 Product Usage: The Cardioplegia Delivery Sets are indicated for delivery of cardioplegic solutions to the heart during cardiopulmonary bypass surgery. | Clear Cardioplegia Delivery Sets | 2 | 10/31/2014 | 10/31/2014 | 11/06/2014 | 1000149028 | Terumo Cardiovascular Systems Corp | Defect in the luer connector can compromise the seal on the female side of the connector. and the compromised seal can introduce air in the table line
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93576 | Z-2614-2010 | Cardiovascular Procedure Kit X-Coated Pump and Table Pack, Item #71552 |  | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93578 | Z-2616-2010 | Cardiovascular Procedure Kit Custom 4:1 Myocardial Set, (MP4 w/ Coil Component Utilized), Item # 63485 | Cardiovascular Procedure Kit***Terumo Cardiovascular Systems Corporation | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93577 | Z-2615-2010 | Cardiovascular Procedure Kit, (MP4 w/ Conducer Component Utilized), labeled in part ***RWJ Hospital***, Item # 71305-03 | Cardiovascular Procedure Kit***Terumo Cardiovascular Systems Corporation | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93573 | Z-2612-2010 | Sarns Cardioplegia Set with w/ PVC Coil from MP4, Item # 15501 | MYO Cardioplegia Delivery Sets | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93584 | Z-2617-2010 | Cardiovascular Procedure Kit X-Coated H/L pack with Cardioplegia, Item # 70067-03 | Cardiovascular Procedure Kit***Terumo Cardiovascular Systems Corporation | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93574 | Z-2613-2010 | Sarns Cardioplegia Set CPLG BLD MP4 4:1 No recirculation, Item # 16010 | MYO Cardioplegia Delivery Sets | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93572 | Z-2611-2010 | Sarns Cardioplegia Set with Conducer Heat Exchanger Cplg/MP4 4:1, Item # 165720 | MYO Cardioplegia Delivery Sets | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=93571 | Z-2610-2010 | Sarns Cardioplegia Set Conducer Heat Exchanger Cplg/MP4 4:1 Bridge, Item# 6375 | MYO Cardioplegia Delivery Sets | 2 | 09/30/2010 | 09/29/2010 | 02/29/2012 | 3000204839 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid.
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92559 | Z-2163-2010 | Terumo Cardiovascular Procedure Kit Custom CONDUCER/MP4 8:1 NO RECIRC. DOUBLE SPIKE P/N: 71032 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92558 | Z-2162-2010 | Terumo Cardiovascular Procedure Kit Custom X-COATED HEART - LUNG PACK P/N: 70608-01 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92498 | Z-2157-2010 | Terumo Cardiovascular Procedure Kit Custom X-COATED PERFUSION PACK P/N 71934 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92482 | Z-2150-2010 | Terumo Cardiovascular Procedure Kit Custom 1:1 CONDUCER/MP4 CRYSTALLOID CARDIOPLEGIA P/N 65246 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92504 | Z-2159-2010 | Terumo Cardiovascular Procedure Kit Custom X-COATED MASTER PACK P/N 72790 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92497 | Z-2156-2010 | Terumo Cardiovascular Procedure Kit Custom X-COATED HEART/LUNG FRAMEPACK P/N 71766PP | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92487 | Z-2151-2010 | Terumo Cardiovascular Procedure Kit Custom MP4, 4:1, PRELIEF, W/RECIRC, COND / PUMP & TABLE PACK P/N 65503/70881-02 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92507 | Z-2161-2010 | Terumo Cardiovascular Procedure Kit Custom ADULT PACK P/N: 131351-02 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92506 | Z-2160-2010 | Terumo Cardiovascular Procedure Kit Custom X-COATED ADULT PACK P/N 72846 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92494 | Z-2154-2010 | Terumo Cardiovascular Procedure Kit Custom ADULT PACK P/N 71546-01 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92492 | Z-2153-2010 | Terumo Cardiovascular Procedure Kit Custom X-COATED ADULT PRESTIGE PACK P/N 71544-01 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92490 | Z-2152-2010 | Terumo Cardiovascular Procedure Kit Custom 8:1 CARDIOPLEGIA SET P/N 70608-01 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92501 | Z-2158-2010 | Terumo Cardiovascular Procedure Kit Custom X-COATED PUMP AND TABLE PACK W/8:1 CARDIOPLEGIA P/N 72642 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92474 | Z-2149-2010 | Terumo Cardiovascular Procedure Kit Custom Open Perfusion Pack P/N 64937 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92495 | Z-2155-2010 | Terumo Cardiovascular Procedure Kit Custom PERFUSION TUBING PACK P/N 71733 | Terumo Custom Kit | 2 | 08/02/2010 | 08/02/2010 | 05/07/2012 | 1000149028 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=200678 | Z-1641-2023 | QUADROX-i Pediatric, Oxygenator used during cardiopulmonary bypass Model Item No. HMO 30000-USA 701070384 HMO 31000-USA 701070388 | QUADROX-i Pediatric | 2 | 05/26/2023 | 05/26/2023 |  | 3008355164 | Maquet Medical Systems USA | 1. An accessory may be improperly placed during packaging, causing creases and damage to the accessory's sterile pouch. 2. Product packaging may exhibit holes, cracks, dents, and crushed areas. Both issues may compromise the sterile barrier of the packaging, and exposure to non-sterile product may result in inflammation, infection, sepsis, and ischemia.
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=180728 | Z-2107-2020 | Ref 801168, MPS Top Cover Replacement Kit, 1 UNITS, UDI: (01)2 0634624 81168 7 (11)190916 (17)290913 (10)059135 (The device is a component of the MPS Console used as replacements for worn or damaged MPS Console top covers) | MPS2 Myocardial Protection System Console | 2 | 05/22/2020 |  | 04/17/2024 | 3001665800 | Quest Medical, Inc. | Additive and Arrest Agent Labels on the MPS Top Cover Replacement Kit, a replacement for worn or damaged MPS Console Top Covers for cardiopulmonary bypass heat exchanger, used to deliver whole blood to the heart during open heart surgery, were swapped and incorrectly placed. This could lead to over delivery of potassium by the user requiring medical intervention.
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=95458 | Z-1924-2011 | Quest MPS and MPS 2 Consoles, Quest Medical, Inc., An Atrion company. Single software controlled device that incorporates a pump, a heater/cooler, temperature monitoring, pressure monitoring, a heat exchanger, an arrest agent pump and an additive pump. The MPS console is used with the MPS sterile disposables with a built in heat exchanger. The device is used with on-pump or off-pump cardiac surgery. | MPS and MPS 2 Consoles | 3 | 04/08/2011 | 04/08/2011 | 04/26/2011 | 3001665800 | Quest Medical, Inc | User update regarding installation of external heater-Cooler units with MPS and MPS2 consoles.
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=25116 | Z-0298-03 | Heat Exchanger Coil (code Number 0006-0011) used in the ''Custom Cardiovascular Procedure Kits'' . | Terumo Cardiovacular Systems Heat Exchanger Coil | 2 | 04/02/2003 | 04/03/2003 | 08/06/2015 | 1000149028 | Terumo Cardiovascular Systems Corp | Heat exchanger may contain loose particulate matter
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=26348 | Z-0664-03 | Sarns MP4 Cardioplegia Sets; catalog number 15501. | Sarns MP4 Cardioplegia Sets | 2 | 03/19/2003 | 03/20/2003 | 06/06/2003 | 1828100 | Terumo Cardiovascular Systems Corp | There is a potential for delivery of a lower than intended amount of cardioplegia solution to the heart during surgery.
2 | DTR | HEAT-EXCHANGER, CARDIOPULMONARY BYPASS | 870.4240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=125652 | Z-1165-2014 | MAQUET PLEGIOX Cardioplegia Heat Exchanger The PLEGIOX Cardioplegia Heat Exchanger is used as part of a cardioplegia system to set and maintain temperature for given flow rates and within the given temperature range of blood cardioplegic and crystalloid cardioplegic solutions during extracorporeal circulation. The product is designed for single use only, for an application period of no longer than 6 hours. | MAQUET | 2 | 03/10/2014 | 03/10/2014 | 04/24/2015 | 3008355164 | Maquet Cardiovascular Us Sales, Llc | The MAQUET post-marketing vigilance program has identified a discrepancy that the desired cardioplegia temperature may not be obtained when the PLEGIOX Cardioplegia Heat Exchanger (PLEGIOX) is used in conjunction with a heater/cooler unit.
2
2 | GAL | SUTURE, ABSORBABLE, NATURAL | 878.4830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=137511 | Z-1839-2015 | Ethicon Fast Absorbing Plain Gut - MultiPass Needles, Sterile, Absorbable Surgical Suture, 5-0 (1.5 metric), 18" (45 cm), Conventional Cutting, Lot #JD6163, Exp: July 2019. Fast absorbing surgical gut suture is a strand of collagenous material prepared from the submucosal layers of the small intestine of healthy sheep, or from the serosal layers of the small intestine of healthy cattle. Fast absorbing surgical gut sutures are sterile and elecit only a slight to minimal tissue reaction during absorption. | Ethicon Fast Absorbing Plain Gut - MultiPass Needles, Sterile, Absorbable Surgical Suture | 2 | 06/25/2015 | 06/25/2015 | 06/06/2017 | 2210968 | Ethicon, Inc. | Ethicon Fast Absorbing Surgical Gut (Plain) Suture is being recalled because the labeling on the box and foil are incomplete.
2 | GAL | SUTURE, ABSORBABLE, NATURAL | 878.4830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=199085 | Z-1388-2023 | Chromic Gut Absorbable Sutures, SURGIGUT* 0 CHR GS-25 90CM X36 (CG-904), SURGIGUT* 1 CHR GS-24 90CM X36 (CG-915), CHROMIC GUT 4-0 75CM V20 (GG-121), CHROMIC GUT 3-0 75CM V20 X36 (GG-122), CHROMIC GUT 2-0 75CM V20 X36 (GG-123), CHROMIC GUT 2-0 75CM V26 (GG-127), CHROMIC GUT 4-0 45CM P12 X36 (SG-5637), CHROMIC GUT 5-0 45CM P-13 X12 (SG-5687G), CHROMIC GUT 5-0 75CM CV23 (UG-202), CHROMIC GUT 4-0 75CM CV23 X36 (UG-203) | Chromic Gut Absorbable Sutures | 2 | 04/12/2023 | 04/12/2023 |  | 1219930 | Covidien, LP | Medtronic was made aware of a supplier calibration issue during the contract sterilization process that may lead to unsterile product. Use of affected product may lead to an unspecified infection or, in the case of sutures used in the eye, pose a loss of vision risk.
2 | GAL | SUTURE, ABSORBABLE, NATURAL | 878.4830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=199087 | Z-1390-2023 | Plain Gut Absorbable Sutures, PLAIN GUT 5/0 18 P-13 X12 (SG-5686), PLAIN GUT 4-0 18 UD SC-11 DA (SG828-2) | Plain Gut Absorbable Sutures | 2 | 04/12/2023 | 04/12/2023 |  | 1219930 | Covidien, LP | Medtronic was made aware of a supplier calibration issue during the contract sterilization process that may lead to unsterile product. Use of affected product may lead to an unspecified infection or, in the case of sutures used in the eye, pose a loss of vision risk.
2 | GAL | SUTURE, ABSORBABLE, NATURAL | 878.4830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=199086 | Z-1389-2023 | Mild Chromic Gut Absorbable Sutures, CH GUT 5-0 12 UD SS-14 SS-14 (G-1766-K), CHROMIC GUT 6-0 45CM HE-1 X12 (G-1780K) | Mild Chromic Gut Absorbable Sutures | 2 | 04/12/2023 | 04/12/2023 |  | 1219930 | Covidien, LP | Medtronic was made aware of a supplier calibration issue during the contract sterilization process that may lead to unsterile product. Use of affected product may lead to an unspecified infection or, in the case of sutures used in the eye, pose a loss of vision risk.
2 | GAL | SUTURE, ABSORBABLE, NATURAL | 878.4830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=171120 | Z-1027-2019 | Plain and Chromic Gut Absorbable Surgical Sutures, sterile. 12 sutures per box. Product Usage: Plain Gut Sutures are indicated for use in general soft tissue approximation and/or ligation, including use in ophthalmic procedures, but not for use in cardiovascular and neural tissues. | Plain gut sutures | 2 | 03/19/2019 |  |  | 3012164473 | CP Medical Inc | Potential compromise of product sterility after routine monitoring for Endotoxins showed an out of limit test result.
2
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=211027 | Z-0744-2025 | TEMA Elbow - Axle Small REF 1590.15.010 TEMA Elbow - Axle Large REF 1590.15.020 TEMA Elbow System is a cemented total elbow prosthesis. | TEMA ELBOW - AXLE | 2 | 12/20/2024 | 12/20/2024 |  | 3008021110 | Limacorporate S.p.A | Due to a potential difference in mechanical behavior between the implanted device and tool(s) used during revision surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=102625 | Z-3082-2011 | Latitude", Tige Humerale, Humeral Stem, Medium 77 mm, Humeral Stem Right, For Cemented use only, Sterile R, REF 0030402, Tornier In. Edina, MN 55435 USA. The Tornier In. Latitude Elbow Prosthesis is intended for total elbow arthroplasty. Prosthetic replacement with this device may be indicated to relieve severe pain or significant disability following the effects of primary or secondary osteoarthritis and rheumatoid arthritis; correction of functional deformities; revision procedures where other treatments or devices have failed; treatment of fractures that are unmanageable using others techniques. The Tornier In. Latitude Elbow Prosthesis is intended for cemented use only. | Tornier Latitude Elbow Prosthesis | 2 | 08/29/2011 | 08/29/2011 | 07/17/2012 | 3000718467 | Tornier, Inc | Tornier has discovered that some models of the Tornier latitude Elbow Prosthesis (Small and Medium Humeral stems) were manufactured with a humeral screw that may not function as designed.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138751 | Z-2428-2015 | Radio-Capitellum, Capitellum, Left, Large, Do Not Reuse, Rx Only. Indicated for use in the elbow for reduction or relief of pain and/or improved elbow function in skeletally mature patients. | Stryker Radio-Capitellum, Capitellum | 2 | 08/20/2015 | 08/20/2015 | 04/29/2016 | 3003070421 | Stryker Howmedica Osteonics Corp. | Stryker is recalling rHead, uHead, Sigmoid Notch, Remotion, Radio Capitellum because packaging integrity (sterile barrier) of the packaging type KIT I may potentially be compromised by transportation.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138752 | Z-2429-2015 | Recon, Radial Implant, Standard Stem, Size 1, Sterile, Do Not Reuse, Rx Only. Intended for replacement of the proximal end of the radius. | Stryker Recon, Radial Implant, Standard Stem | 2 | 08/20/2015 | 08/20/2015 | 04/29/2016 | 3003070421 | Stryker Howmedica Osteonics Corp. | Stryker is recalling rHead, uHead, Sigmoid Notch, Remotion, Radio Capitellum because packaging integrity (sterile barrier) of the packaging type KIT I may potentially be compromised by transportation.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=138723 | Z-2426-2015 | rHead Lateral Stem, Size 1, Do Not Reuse, Sterile. Intended for replacement of the proximal end of the radius. | Stryker rHead Lateral Stem | 2 | 08/20/2015 | 08/20/2015 | 04/29/2016 | 3003070421 | Stryker Howmedica Osteonics Corp. | Stryker is recalling rHead, uHead, Sigmoid Notch, Remotion, Radio Capitellum because packaging integrity (sterile barrier) of the packaging type KIT I may potentially be compromised by transportation.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161950 | Z-2695-2018 | Radio Capitellum Large, Right. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161955 | Z-2700-2018 | Radio Capitellum, Recon Head, Size #2. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161960 | Z-2705-2018 | rHead Standard Extended Stem, 6mm Collar, Size 3. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161954 | Z-2699-2018 | Radial Head Assy Size #4, Rad. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161958 | Z-2703-2018 | rHead Standard Extended Stem, 6mm Collar, Size 1. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161959 | Z-2704-2018 | rHead Standard Extended Stem, 6mm Collar, Size 2. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161961 | Z-2706-2018 | rHead Standard Extended Stem, 6mm Collar, Size 4. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161963 | Z-2708-2018 | rHead Recon Extended Stem Size 2. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161965 | Z-2710-2018 | Radio Capitellum Broach Large, Right. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161948 | Z-2693-2018 | Radio Capitellum Large, Left Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161956 | Z-2701-2018 | Radio Capitellum, Recon Head, Size #3. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161953 | Z-2698-2018 | Radial Head Assy Size #3, Rad. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161951 | Z-2696-2018 | Radio Capitellum Small, Right. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161962 | Z-2707-2018 | rHead Recon Extended Stem Size 1. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161949 | Z-2694-2018 | Radio Capitellum Small, Left. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161964 | Z-2709-2018 | rHead Recon Extended Stem Size 3. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161957 | Z-2702-2018 | Radio Capitellum, Recon Head, Size #4. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161952 | Z-2697-2018 | Radial Head Assy Size #2, Rad. Orthopedic implant for replacement of the proximal end of the radius in the Stryker rHead Radial Head and Uni-Elbow system. |  | 2 | 2018-08-08 |  | 05/18/2023 | 3003070421 | Howmedica Osteonics Corp. | Revision rate identified as higher than other proximal radial head replacement devices based on literature review. Product may not perform at expected level post-surgery.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=213028 | Z-1651-2025 | stryker Latitude EV HUMERAL REVISION STEM TRIAL, L: 150mm, Size: LARGE, Side: RIGHT, REF 9030103; Prosthesis, Elbow, Semi-Constrained, Cemented | stryker | 2 | 04/22/2025 | 04/22/2025 |  | 3000718467 | Tornier, Inc | A specific lot of Latitude Humeral Trial Stems has the incorrect direct part marking. " Laser Marked: 9030103, Humeral Stem Trial 150mm Large / Right, Yellow color indicator (incorrect) " Device Actually: 9030092, Humeral Stem Trial 150mm Medium / Left, Red color indicator (correct)
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=43150 | Z-0531-06 | Depuy brand Acclaim Elbow linked ulnar bearing assembly (ulnar bearing and locking screw), sterile; catalog # 1722-63-000. | Acclaim | 2 | 02/15/2006 | 02/16/2006 | 06/16/2006 | 1818910 | Depuy Orthopaedics, Inc. | The ulnar bearing may damage the implant's polyethylene sleeve such that revision surgery is necessary.
2 | JDB | PROSTHESIS, ELBOW, SEMI CONSTRAINED, CEMENTED | 888.3160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=42917 | Z-0530-06 | Depuy brand Acclaim Elbow linked ulnar bearing, sterile; catalog # 1722-60-000. | Acclaim | 2 | 02/15/2006 | 02/16/2006 | 06/16/2006 | 1818910 | Depuy Orthopaedics, Inc. | The ulnar bearing may damage the implant's polyethylene sleeve such that revision surgery is necessary.
2
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=114542 | Z-0537-2013 | COONMAD/MORREY TOTAL ELBOW ELBOW COMPONENT REPLACEMENT SET BUSHINGS & PINS FOR USE W/SIZE REGULAR ELBOW ELBOW COMP REPLACEMENT-REG SET STERILE PROD CLASS: 4100 OPERATION:9000 QTY/PKG:1 Product Usage: The ulnar and humeral components are connected by placing the hollow, outer axis pin across the two components and securing It with the solid Internal axis pin. The Inner and outer pin connection Is articulation point of the placement elbow. | Coonmad/Morrey | 2 | 12/17/2012 | 12/18/2012 | 04/17/2014 | 1000220733 | Zimmer, Inc. | Zimmer is initiating a lot specific recall because affected lots could include an extra Small Inner Pin instead of the required Regular Inner Pin in the Coonrad/Morrey Elbow Replacement Set and the New Ulnar Revision Kit .
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=114543 | Z-0538-2013 | Coonrad/Morrey New Ulnar Revision Kit for Use with Non-Interchangeable Regular Humerals STERILE Rx, Tivanium QTY/PKG:1 Product Usage: The ulnar and humeral components are connected by placing the hollow, outer axis pin across the two components and securing It with the solid Internal axis pin. The Inner and outer pin connection Is articulation point of the placement elbow. | Coonrad/Morrey | 2 | 12/17/2012 | 12/18/2012 | 04/17/2014 | 1000220733 | Zimmer, Inc. | Zimmer is initiating a lot specific recall because affected lots could include an extra Small Inner Pin instead of the required Regular Inner Pin in the Coonrad/Morrey Elbow Replacement Set and the New Ulnar Revision Kit .
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=176817 | Z-0301-2020 | Coorad Morrey Total Elbow, Item Nos. 32810502501 32810502701 32810502900 32810503000 32810600012 32810600013 32855510796 32855510797 32855510799 32855510800 32855510801 32855510805 32855510822 32855510843 32855510844 | Coorad Morrey Total Elbow | 2 | 11/07/2019 |  | 05/19/2020 | 1000220733 | Zimmer Biomet, Inc. | Non-highly polished implants may adhere to the LDPE bag (previous packaging configuration).
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=73486 | Z-0032-2009 | Biomet brand Discovery Elbow, 3 mm x 75 mm right ulna with bearing/bond coat, TI 6AL 4V Alloy/ARCOM UHMWPE, For cemented use only, sterile, Biomet Orthopedics, Inc., Warsaw, IN; REF 114813. The product intended use is for Orthopedic implant. | Biomet brand Discovery Elbow | 2 | 10/03/2008 | 10/03/2008 | 09/22/2009 | 1825034 | Biomet, Inc. | The component in the package is not the correct size.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=73487 | Z-0033-2009 | Biomet brand Discovery Elbow, 3 mm x 115 mm right ulna with bearing/bond coat, TI 6AL 4V Alloy/ARCOM UHMWPE, For cemented use only, sterile, Biomet Orthopedics, Inc., Warsaw, IN; REF 114817. The product intended use is for Orthopedic implant. | Biomet brand Discovery Elbow | 2 | 10/03/2008 | 10/03/2008 | 09/22/2009 | 1825034 | Biomet, Inc. | The component in the package is not the correct size.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=221858 | Z-3195-2026 | AXLE PIN ASSY LARGE ELBOW COBALT CHROME Catalog No. C-M112-2-500 Orthopaedic implant components manufactured from UHMWPE (polyethylene) | Stryker | 2 | 09/16/2026 | 09/16/2026 |  | 3003070421 | Howmedica Osteonics Corp. | Affected devices were manufactured with Ultra High Molecular Weight Polyethylene (UHMWPE) raw material aged in excess of five years. Product has the potential for elevated levels of oxidation, which may result in increased wear, leading to hazards such as excessive wear debris, material fragments, fractured device, or disassociated device. Potential harms include revision surgery, pain, and inflammation.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=149116 | Z-2752-2016 | Coonrad/ Morrey Elbow Pin/ Bushing Replacement size Kit Extra Small. This product is sold sterile. The parts are placed into an inner cavity and are sealed using a nitrogen process. The inner cavity is then placed into an outer cavity and heat sealed with outer Tyvek lid on the outer cavity. The patient record label is placed on the Tyvek outerlid. The sealed cavities and package insert is placed into a folding carton, sealed and shrink wrapped. Total elbow prosthesis designed for use with bone cement | Coonrad/ Morrey Elbow Pin/ Bushing Replacement size Kit Extra Small | 2 | 09/08/2016 | 09/08/2016 | 06/12/2017 | 1000220733 | Zimmer Biomet, Inc. | Zimmer Biomet is conducting a voluntary medical device recall of a single lot of the Coonrad/ Morrey Elbow Pin/ Bushing Replacement size Kit Extra Small because the lot was packaged with a missing humeral bushing. The missing bushing could result in a surgical delay or additional surgery if a replacement is not available.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=127549 | Z-1925-2014 | DISCOVERY ELBOW prosthesis, 4 MM X 100 MM RIGHT HUMERAL FLANGED I BOND COAT Tl 6AL 4V ALLOY. Elbow joint replacement prostheses are intended for primary and revision joint arthroplasty for use in cemented applications. | DISCOVERY ELBOW | 2 | 06/26/2014 | 06/26/2014 | 03/16/2015 | 1825034 | Biomet, Inc. | Surface finish is different than specified. The implant has 30 grit blast on the exterior of the flange when the print calls for the glass bead blast. The patient may experience pain requiring surgical intervention.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=100452 | Z-2660-2011 | Condyle Kit with Hexalobular, REF 114700 DISCOVERY ELBOW SYSTEM, HUMERAL CONDYLE SET - HEXALOBULAR, CO-CR-MO / YI-GAL-4V ALLOY, IF USED FOR REVISION SURGERY, INSTRUMENTS 414926 & 414923 MAY BE REQUIRED LOT 183300, 1 SET, BIOMET ORTHOPEDICS 56 EAST BELL DRIVE, P.O. BOX 587 WARSAW, IN 46581 USA EXPIRY DATE: 2021-03 Elbow joint replacement prosthesis intended for primary and revision joint arthroplasty for use in cemented applications. | Discovery Elbow System with Humeral Condyle | 2 | 06/24/2011 | 06/24/2011 | 08/29/2011 | 1825034 | Biomet, Inc. | The firm initiated this recall after becoming aware that a Discovery Condyle Kit with Hexalobula contained two male condyles instead of a male and female condyle.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=32155 | Z-0767-04 | Zimmer brand Coonrad/Morrey total elbow interchangeable ulnar assembly, small, left, 3 in. length. tivanium TI-6AL-4V alloy/UHMWPE, sterile, for use with bone cement. Cat. no. 32-8105-53-01. | Coonrad/Morrey total elbow. | 2 | 04/05/2004 | 07/20/2004 | 06/26/2004 | 1000220733 | Zimmer Inc. | Right ulna implants were etched and labeled as left ulna implants. Left ulna implants may have been etched and labeled as right ulna implants.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=32157 | Z-0768-04 | Zimmer brand Coonrad/Morrey total elbow interchangeable ulnar assembly, small, right, 3 in. length. tivanium TI-6AL-4V alloy/UHMWPE, sterile, for use with bone cement. Cat. no. 32-8105-53-02. | Coonrad/Morrey total elbow. | 2 | 04/05/2004 | 07/20/2004 | 06/26/2004 | 1000220733 | Zimmer Inc. | right ulna implants were etched and labeled as left ulna implants. Left ulna implants may have been etched and labeled as right ulna implants.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=77694 | Z-1057-2009 | Zimmer Coonrad/Morrey Total Elbow, Interchangeable Ulnar Assembly, plasma sprayed, extra small, left, long, sterile, Zimmer, Warsaw, IN; REF 32-8105-093-01. The device is utilized for a total elbow arthroplasty surgical procedure. | Zimmer Coonrad/Morrey Total Elbow | 2 | 04/02/2009 | 04/02/2009 | 12/22/2009 | 1000220733 | Zimmer Inc. | The surgical technique is inadequate for extra-small sized implants, resulting in the possibility of the provisional becoming stuck in the bone.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=77692 | Z-1055-2009 | Zimmer Coonrad/Morrey Total Elbow, Interchangeable Ulnar Assembly, plasma sprayed, extra small, left, 3 inch length, sterile, Zimmer, Warsaw, IN; REF 32-8105-043-01. The device is utilized for a total elbow arthroplasty surgical procedure. | Zimmer Coonrad/Morrey Total Elbow | 2 | 04/02/2009 | 04/02/2009 | 12/22/2009 | 1000220733 | Zimmer Inc. | The surgical technique is inadequate for extra-small sized implants, resulting in the possibility of the provisional becoming stuck in the bone.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=77693 | Z-1056-2009 | Zimmer Coonrad/Morrey Total Elbow, Interchangeable Ulnar Assembly, plasma sprayed, extra small, right, 3 inch length, sterile, Zimmer, Warsaw, IN; REF 32-8105-043-02. The device is utilized for a total elbow arthroplasty surgical procedure. | Zimmer Coonrad/Morrey Total Elbow | 2 | 04/02/2009 | 04/02/2009 | 12/22/2009 | 1000220733 | Zimmer Inc. | The surgical technique is inadequate for extra-small sized implants, resulting in the possibility of the provisional becoming stuck in the bone.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=77691 | Z-1054-2009 | Zimmer Coonrad/Morrey Total Elbow, Interchangeable Humeral Assembly, extra small, 6 inch length, sterile, Zimmer Inc, Warsaw, IN; REF 32-8105-027-06. The device is utilized for a total elbow arthroplasty surgical procedure. | Zimmer Coonrad/Morrey Total Elbow | 2 | 04/02/2009 | 04/02/2009 | 12/22/2009 | 1000220733 | Zimmer Inc. | The surgical technique is inadequate for extra-small sized implants, resulting in the possibility of the provisional becoming stuck in the bone.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=77695 | Z-1058-2009 | Zimmer Coonrad/Morrey Total Elbow, Interchangeable Ulnar Assembly, plasma sprayed, extra small, right, long, sterile, Zimmer, Warsaw, IN; REF 32-8105-093-02. The device is utilized for a total elbow arthroplasty surgical procedure. | Zimmer Coonrad/Morrey Total Elbow | 2 | 04/02/2009 | 04/02/2009 | 12/22/2009 | 1000220733 | Zimmer Inc. | The surgical technique is inadequate for extra-small sized implants, resulting in the possibility of the provisional becoming stuck in the bone.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=77690 | Z-1053-2009 | Zimmer Coonrad/Morrey Total Elbow, Interchangeable Humeral Assembly, extra small, 4 inch length, sterile, Zimmer Inc., Warsaw, IN; REF 32-8105-027-04. The device is utilized for a total elbow arthroplasty surgical procedure. | Zimmer Coonrad/Morrey Total Elbow | 2 | 04/02/2009 | 04/02/2009 | 12/22/2009 | 1000220733 | Zimmer Inc. | The surgical technique is inadequate for extra-small sized implants, resulting in the possibility of the provisional becoming stuck in the bone.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=162220 | Z-1239-2018 | Coonrad/Morrey Elbow Implant System humeral assembly 4IN XSML, Item Number: 32-8105-027-04 (UDI: (01)0088902427 4006(17)220731(10)63694912) | Coonrad/Morrey Elbow Implant System | 2 | 03/28/2018 | 03/14/2018 | 04/09/2019 | 1000220733 | Zimmer Biomet, Inc. | One lot of humeral stem implants do not include the humeral bushing components.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=198354 | Z-1261-2023 | TEMA Axle Small, REF 1590.15.010 (additional information is pending receipt of the firm's product label/labeling) | TEMA Axle Small | 2 | 03/16/2023 | 03/16/2023 |  | 3008021110 | Limacorporate S.p.A | Due to the incorrect component, that is not cleared to be marketed in the U.S., being included in its elbow prosthesis.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=106933 | Z-1071-2012 | Discovery Elbow System Humeral Condyle Set- Hexalobular, REF 114700, Lot 797550 Alloy, Sterile, Biomet Orthopedics Warsaw IN 46581. Intended for primary and revision joint arthroplasty for use in cemented applications. | Discovery Elbow System Humeral Condyle Set- Hexalobular | 2 | 02/21/2012 | 02/21/2012 | 12/12/2012 | 1825034 | Biomet, Inc. | A Discovery Condyle Kit w/ Hexalobula was opened during an elbow joint replacement surgery, and was found to contain two male condyles.
2 | JDC | PROSTHESIS, ELBOW, CONSTRAINED, CEMENTED | 888.3150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=178851 | Z-1000-2020 | Coonrad/Morrey Total Elbow Humeral Assembly, Extra Small, 4-inch Length. Model No. 32-8105-027-04. UDI (01)00889024274006 (17)241031(10)64481139 The device is a Humeral component; elbow joint metal/polymer constrained cemented prosthesis. | Coonrad/Morrey Total Elbow Humeral Assembly | 2 | 02/05/2020 |  | 01/13/2021 | 1000220733 | Zimmer Biomet, Inc. | The lot contained two outer pins instead of one outer and one inner pin.
2
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=122646 | Z-0404-2014 | Synthes 14.5 mm Outer Protection Sleeve for Suprapatellar Synthes 14.5 mm Outer Protection Sleeve for Suprapatellar is used as part of the Synthes Tibial Nail System EX to stabilize fractures of the proximal and distal tibia and the tibia shaft, open and closed tibia shaft fractures, certain pre- and post- isthmic fractures, and tibial malunions and non-unions. | Synthes 14.5 mm Outer Protection Sleeve for Suprapatellar | 2 | 11/26/2013 | 11/26/2013 | 08/18/2015 | 3008812560 | Synthes (USA) Products LLC | Complaints were received reporting that a product packaged as a 14.5 mm Outer Protection Sleeve for Suprapatellar contained a 12.0 mm Outer Protection Sleeve for Suprapatellar.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=122363 | Z-0407-2014 | Synthes Titanium Cannulated Humeral Nail Intended to aid in the alignment and stabilization of humeral fractures | Synthes Titanium Cannulated Humeral Nail | 2 | 11/26/2013 | 11/27/2013 | 09/10/2015 | 3005180112 | Synthes USA HQ, Inc. | The packaging incorrectly depicts the Synthes Titanium Cannulated Humeral Nail as a 3 distal hole nail instead of a 2 distal locking hole nail.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=159397 | Z-0097-2018 | smith&nephew TRIGEN INTERTAN 10S, REF 71675384, 10 MM X 18 CM NAIL, 130o (degrees), SET SCREW PREASSEMBLED IN NAIL | TRIGEN INTERTAN 10S | 2 | 11/16/2017 |  |  | 1020279 | Smith & Nephew, Inc. | A single lot of INTERTAN 10S 10 MM X 18 CM 130 D was manufactured with the screw inserted upside down.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=216349 | Z-0229-2026 | LOW PROF LCKNG SCREW, F/IM NAIL 5.0MM / L, 42MM / XL25/ STER. Product Code: 04.045.342TS. Used in RFN-ADVANCEDTM Retrograde Femoral Nailing (RFNA) System and the Tibial Nail Advanced System. | Locking Screw | 2 | 10/21/2025 | 10/21/2025 |  | 3008812560 | Synthes (USA) Products LLC | Although the etched length on the screw is correct, the length marked on the external packaging is incorrect. Specifically, 42mm screws were packaged and labeled as 46mm, while 46mm screws were packaged and labeled as 42mm.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=216348 | Z-0228-2026 | LOW PROF LCKNG SCREW, F/IM NAIL 5.0MM / L, 46MM / XL25/ STER. Product Code: 04.045.346TS. Used in RFN-ADVANCEDTM Retrograde Femoral Nailing (RFNA) System and the Tibial Nail Advanced System. | Locking Screw | 2 | 10/21/2025 | 10/21/2025 |  | 3008812560 | Synthes (USA) Products LLC | Although the etched length on the screw is correct, the length marked on the external packaging is incorrect. Specifically, 42mm screws were packaged and labeled as 46mm, while 46mm screws were packaged and labeled as 42mm.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=40870 | Z-0012-06 | EBI Trochanteric 11mm Long Nail. |  | 2 | 10/06/2005 | 10/07/2005 | 07/03/2006 | 2242816 | EBI, L.P. | Upon evaluation, the firm found that the fatigue strength of these nails was below the required value as specified in the design specification.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=128687 | Z-2258-2014 | TRIGEN(R) INTERTAN(TM) 10MM X 18 CM NAIL, 125o, REF 71675201, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 11.5MM X 18 CM NAIL, 125o, REF 71675202, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 13MM X 18 CM NAIL, 125o, REF 71675203, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 10MM X 20 CM NAIL, 125o, REF 71675204, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 11.5MM X 20 CM NAIL, 125o, REF 71675205, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 13MM X 20 CM NAIL, 125o, REF 71675206, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 10MM X 18 CM NAIL, 130o, REF 71675207, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 11.5MM X 18 CM NAIL, 130o, REF 71675208, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 13MM X 18 CM NAIL, 130o, REF 71675209, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 10MM X 20 CM NAIL, 130o, REF 71675210, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 11.5MM X 20 CM NAIL, 130o, REF 71675211, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 13MM X 20 CM NAIL, 130o, REF 71675212, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 10MM X 26 CM NAIL, 125o, LEFT, REF 71675213, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 10MM X 32 CM NAIL, 125o, LEFT, REF 71675219, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 10MM X 34 CM NAIL, 125o, LEFT, REF 71675221, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 10MM X 34 CM NAIL, 125o, RIGHT, REF 71675222, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 10MM X 36 CM NAIL, LEFT, 125o, REF 71675223, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 10MM X 38 CM NAIL, 125o, LEFT, REF 71675225, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 10MM X 40 CM NAIL, 125o, LEFT, REF 71675227, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 10MM X 40 CM NAIL, 125o, RIGHT, REF 71675228, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 11.5MM X 28 CM NAIL, 125o, RIGHT, REF 71675238, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 11.5MM X 32 CM NAIL, 125o, LEFT, REF 71675241, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 11.5MM X 34 CM NAIL, 125o, LEFT, REF 71675243, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 11.5MM X 34 CM NAIL, 125o, RIGHT, REF 71675244, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 11.5MM X 36 CM NAIL, 125o, LEFT, REF 71675245, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 11.5MM X 36 CM NAIL, 125o, RIGHT, REF 71675246, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 11.5MM X 38 CM NAIL, 125o, LEFT, REF 71675247, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 11.5MM X 38 CM NAIL, 125o, RIGHT, REF 71675248, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 11.5MM X 40 CM NAIL, 125o, LEFT, REF 71675249, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 11.5MM X 40 CM NAIL, 125o, RIGHT, REF 71675250, QTY (1), STERILE R, Smith & Nephew, Memphis, Inc., TN 38116 TRIGEN(R) INTERTAN(TM) 10MM X 26 CM NAIL, 130o, LEFT, REF 71675257, QTY (1), STERILE R, Smi | TRIGEN(R) INTERTAN(TM) NAIL | 2 | 08/17/2014 | 08/17/2014 | 08/20/2019 | 1020279 | Smith & Nephew, Inc. | There was an error in manufacturing which resulted in the proximal lag screw hole being positioned incorrectly in several batches of TRIGEN(R); INTERTAN(TM); NAILS.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=182151 | Z-2685-2020 | TRIGEN, Intramedullary nail - Product Usage: is intended to provide resistance for screw back-out. It provides additional friction to help hold the screws in place. | TRIGEN | 2 | 07/22/2020 |  | 04/28/2021 | 1020279 | Smith & Nephew, Inc. | A manufacturing error which resulted in the re-sterilization of the ultra-high-molecular-weight polyethylene (UHMWPE) sleeve.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=137807 | Z-2051-2015 | TITANIUM END CAP WITH T40 STARDRIVE RECESS, 0MM EXTENSION FOR TITANIUM TIBIAL NAILS-EX; Synthes Tibial Nail System EX is intended to stabilize fractures of the proximal and distal tibia and the tibial shaft; open and closed tibial shaft fractures; certain pre-and post-isthmic fractures; and tibial malunions and non-unions. | TITANIUM END CAP WITH T40 STARDRIVE RECESS, 0MM EXTENSION FOR TITANIUM TIBIAL NAILS-EX | 2 | 07/08/2015 | 07/08/2015 | 09/15/2015 | 3005180112 | Synthes, Inc. | A specific part and lot number was packaged and shipped prior to the completion of a required internal inspection.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=80216 | Z-1349-2009 | Stryker, Gamma3 Plus Obese Targeting System, Operative Technique for Hip Fracture System, Trochanteric and Long Nails. | Gamma3 Plus | 2 | 06/01/2009 | 06/01/2009 | 08/17/2012 | 3003070421 | Stryker Howmedica Osteonics Corp. | Operative Technique: further clarification and guidance is needed to help reduce the potential for distal mis-drilling of the short Gamma3 Trochanteric Nails.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=80218 | Z-1351-2009 | Stryker, Gamma3, The Compact Version of the Gamma Nail System, Operative Technique for Hip Fracture System, Trochanteric and Long Nails. | Gamma3 | 2 | 06/01/2009 | 06/01/2009 | 08/17/2012 | 3003070421 | Stryker Howmedica Osteonics Corp. | Operative Technique: further clarification and guidance is needed to help reduce the potential for distal mis-drilling of the short Gamma3 Trochanteric Nails.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=80217 | Z-1350-2009 | Stryker, Gamma3S Trochanteric Nail 170, Operative Technique for Hip Fracture System. | Gamma3S | 2 | 06/01/2009 | 06/01/2009 | 08/17/2012 | 3003070421 | Stryker Howmedica Osteonics Corp. | Operative Technique: further clarification and guidance is needed to help reduce the potential for distal mis-drilling of the short Gamma3 Trochanteric Nails.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=193510 | Z-1133-2022 | INTERTAN 1.5 NAIL 13mm X 42cm 130DEG. Left INTERTAN 1.5 NAIL 13mm X 42cm 130DEG. Right | TRIGEN InterTAN Nail | 2 | 05/25/2022 |  |  | 3002465897 | Smith & Nephew Orthopaedics GmbH | Right nails were anodized, marked, and labelled as left nails and vice versa
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=126852 | Z-1637-2014 | TRIGEN INTERTAN 1.5 NAIL: (1) REF 71676515, 10 MM X 10 CM, 125 degree, LEFT, (2) REF 71676556, 11.5 MM X 36 CM, 125 degree, RIGHT, (3) REF 71676558, 11.5 MM X 38 CM, 125 degree, RIGHT, (4) REF 71676561, 11.5 MM X 42 CM, 125 degree, LEFT, Smith & Nephew, Inc. Orthopedic fixation device. | TRIGEN INTERTAN Nail | 2 | 05/20/2014 | 05/20/2014 | 07/01/2017 | 1020279 | Smith & Nephew Inc | A manufacturing error resulted in the set screw being inserted too deep in the affected nails.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=126851 | Z-1636-2014 | TRIGEN INTERTAN NAIL: (1) REF 71675201, 10 MM X 18 CM, 125 degree, (2) REF 71675202, 11.5 MM X 18 CM, 125 degree, (3) REF 71675203, 13 MM X 18 CM, 125 degree, (4) REF 71675204, 10 MM X 20 CM, 125 degree, (5) REF 71675205, 11.5 MM X 20 CM, 125 degree, (6) REF 71675207, 10 MM X 18 CM, 130 degree, (7) REF 71675208, 11.5 MM X 18 CM, 130 degree, (8) REF 71675211, 11.5 MM X 20 CM, 130 degree, Smith & Nephew, Inc. Orthopedic fixation device. | TRIGEN INTERTAN NAIL | 2 | 05/20/2014 | 05/20/2014 | 07/01/2017 | 1020279 | Smith & Nephew Inc | A manufacturing error resulted in the set screw being inserted too deep in the affected nails.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90078 | Z-1549-2010 | 7.5mm Ti Solid Humeral Nail 250 mm Sterile Catalog number 462.725S Synthes Titanium Solid Humeral Nail System is intended for intramedullary fixation of a variety of humeral fractures using either an antegrade or retrograde approach. The indications for the Titanium Solid Humeral Nail System includes: diaphyseal fractures of the humeral shaft, fractures of the proximal humerus, proximal humeral fractures with diaphyseal extension, and impending pathological fractures. | 7.5mm Ti Solid Humeral Nail 250 mm Sterile | 2 | 05/07/2010 | 05/07/2010 | 10/20/2014 | 3005180112 | Synthes USA (HQ), Inc. | potential compatibility issue
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90076 | Z-1548-2010 | 7.5mm Ti Solid Humeral Nail 230 mm Sterile Catalog number 462.723S Synthes Titanium Solid Humeral Nail System is intended for intramedullary fixation of a variety of humeral fractures using either an antegrade or retrograde approach. The indications for the Titanium Solid Humeral Nail System includes: diaphyseal fractures of the humeral shaft, fractures of the proximal humerus, proximal humeral fractures with diaphyseal extension, and impending pathological fractures. | 7.5mm Ti Solid Humeral Nail 230 mm Sterile | 2 | 05/07/2010 | 05/07/2010 | 10/20/2014 | 3005180112 | Synthes USA (HQ), Inc. | potential compatibility issue
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=153398 | Z-1858-2017 | various polyethylene implants BIO PLUG 10MM Product Usage: Non-weight bearing accessory. Sterile, optional speciality components that are non-bioabsorbable, implantable devices designed to be inserted into the medullary canal of a bone during orhtopedic surgery to prevent cement progression in the diaphysis and, therefore, facilitate cement pressurization during the introduction of an implant. | various polyethylene implants BIO PLUG | 2 | 04/18/2017 | 03/27/2017 | 01/29/2018 | 1825034 | Zimmer Biomet, Inc. | Endotoxin levels higher than process maximum limits were discovered in the identified polyethylene components.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=115886 | Z-0943-2013 | INTERNAL HEX RECON SCREW, TI-6AL-4V, REF 7162380, TRIGEN (R) INTERNAL HEX CAPTURED RECON SCREW, QTY: (1), 6.4 MM, 80 MM, TITANIUM, Smith & Nephew, Inc. Orthopedic. Bone screw for intramedullary rod fixation. | INTERNAL HEX RECON SCREW | 2 | 03/13/2013 | 03/13/2013 | 01/11/2016 | 1020279 | Smith & Nephew Inc | 75 mm screws were packaged and labeled using an 80 mm screw outer box and label, while 80 mm screws were packaged and labeled using a 75 mm screw outer box and label.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=115887 | Z-0944-2013 | INTERNAL HEX RECON SCREW, TI-6AL-4V, REF 7162375, TRIGEN (R) INTERNAL HEX CAPTURED RECON SCREW, QTY: (1), 6.4 MM, 75 MM, TITANIUM, Smith & Nephew, Inc. Orthopedic. Bone screw for intramedullary rod fixation. | INTERNAL HEX RECON SCREW | 2 | 03/13/2013 | 03/13/2013 | 01/11/2016 | 1020279 | Smith & Nephew Inc | 75 mm screws were packaged and labeled using an 80 mm screw outer box and label, while 80 mm screws were packaged and labeled using a 75 mm screw outer box and label.
2 | JDS | NAIL, FIXATION, BONE | 888.3030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184841 | Z-0772-2021 | Bone fixation nail. TRIGEN INTERTAN - Product Usage: indicated for fractures of the femur including simple shaft fractures, comminuted shaft fractures, spiral shaft fractures, long oblique shaft fractures and segmental shaft fracture; | TRIGEN INTERTAN | 2 | 01/06/2021 |  | 01/06/2023 | 1020279 | Smith & Nephew, Inc. | Manufacturing packaging error.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=184917 | Z-0726-2021 | PORTEX Bivona Cuffless Neonatal/Pediatric FlexTend Tracheostomy Tube, REF 60PFSS45 | Bivona Cuffless Neonatal/Pediatric FlexTend Tracheostomy Tube | 2 | 12/28/2020 |  | 12/20/2022 | 3012307300 | Smiths Medical ASD Inc. | A specific model and lot number of Bivona Cuffless FlexTend Tracheostomy Tube may have been incorrectly packaged with a V-neck flange instead of a straight flange.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=49453 | Z-0260-2007 | Custom 5.5 mm Bivona Uncuffed Hyperflex Tracheostomy Tube, I.D. 5.5 mm, O.D. 8.0 mm, length 4.75 mm, Mfrd by Smiths Medical Critical Care; REF HU05JS55NNA112S. | Bivona Tracheostomy Tube | 3 | 12/06/2006 | 12/07/2006 | 01/12/2007 | 1824231 | Smiths Medical ASD, Inc. | Mislabeled as to size: The product is 47.5 mm in length, but the label lists the length as 4.75 mm.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=64375 | Z-0100-2008 | Portex Perfit Percutaneous Tracheostomy Tube Replacement Kit, Sterile Ref: 536090 | Smiths Portex | 2 | 11/06/2007 | 11/07/2007 | 05/08/2012 | 1217052 | Smiths Medical ASD, Inc. | Mislabeled. The product is labeled as a size 9.0 mm tracheostomy tube, but the package contains a 7.0 mm tracheostomy tube.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=203239 | Z-0205-2024 | Smiths Medical Portex Bivona Pediatric Tracheostomy Tube, Cuffless, Model 60P045, Sterile. | Portex Bivona | 2 | 10/31/2023 | 10/31/2023 |  | 3012307300 | Smiths Medical ASD Inc. | The unit box is labeled with an incorrect expiration date (the manufacturing date of 3/17/2021) but the individual packaging contains the correct expiration date (3/17/2026).
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=149196 | Z-0044-2017 | Willy Rusch GmbH Tracheostomy Tube Set Cannulation of tracheotomised patients, particularly in the case of a narrow tracheostoma and narrow-lumen tracheas. When using cannulas with curr. For short term intermittent inflation. | Rusch TracheoFlex Tracheostomy Tube | 1 | 10/18/2016 | 10/18/2016 | 03/21/2017 | 3005747797 | Teleflex Medical | The connector may disconnect from the tracheostomy tube during use.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=48075 | Z-0046-2007 | Bivona Adult Tracheostomy Tube, Labeled I.D. 6.0 mm, O.D. 8.7 mm, A 7.0 mm, B 0 mm, C 53.0 mm, Mfrd by Smiths Medical Critical Care; REF 60A160. | Bivona | 3 | 10/18/2006 | 10/19/2006 | 12/04/2007 | 1824231 | Smiths Medical ASD, Inc. | The length of the shaft is incorrectly stated as 7 mm, instead of the correct 17 mm, on the outer packaging, although the tray label is correct.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=158327 | Z-3213-2017 | UniPerc(R) Adjustable Flange Tracheostomy Tube kit, Reorder Number 100/897/080 CZ The UniPerc(R) Tracheostomy Tubes can be used during the surgical placement of a stoma or as a replacement tube to an already existing stoma for patients with larger necks. The tracheostomy tubes are removed and may be replaced after no more than 30 days with an identical device. The range may include the following configurations: cuffed or uncuffed tracheostomy tubes, tracheostomy tube inner cannula, cleaning brushes and tracheostomy tube holder. | UniPerc(R) Adjustable Flange Tracheostomy Tube kit | 2 | 09/26/2017 |  | 06/18/2019 | 3012307300 | Smiths Medical ASD Inc. | Smiths Medical became aware that one lot of the 8.0mm UniPerc(R) Adjustable Flange Tracheostomy Tube kit, product reorder number 100/897/080 CZ, lot number 3308971, contains an incorrect 9.0mm sized obturator instead of the correct 8.0mm sized obturator.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34784 | Z-1471-04 | Laryngectomy Tubes, Air Lon Inhalation Set (Inner Cannula with 15 mm Adapter) Nylon, item number 1050200 size 8 and item number 1050202 size 10. | Air Lon Inhalation Set (Inner Cannula with 15 mm Adapter) Nylon Laryngectomy Tubes | 2 | 09/14/2004 | 09/15/2004 | 11/23/2005 | 2511556 | Premier Dental Products Co | foreign objects (polishing stones) in the tubes
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34783 | Z-1470-04 | Laryngectomy Tubes, Air Lon Nylon, item number1050170 size 8 and item number 1050172 size 10. | Air Lon Nylon Laryngectomy Tubes | 2 | 09/14/2004 | 09/15/2004 | 11/23/2005 | 2511556 | Premier Dental Products Co | foreign objects (polishing stones) in the tubes
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34781 | Z-1468-04 | Laryngectomy Tubes, Jackson Original Stainless Steel, item number 1037214 size 4, item number 1037215 size 5, item number 1037216 size 6, item number 1037217 size 7, item number 1037218 size 8, item number 1037219 size 9, and item number 1037220 size 10. | Jackson Original Stainless Steel Laryngectomy Tubes | 2 | 09/14/2004 | 09/15/2004 | 11/23/2005 | 2511556 | Premier Dental Products Co | foreign objects (polishing stones) in the tubes
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34780 | Z-1467-04 | Laryngectomy Tubes, Martin Stainless Steel, item number 1036128 size 8 and item number 1036130 size10. | Martin Stainless Steel Laryngectomy Tubes | 2 | 09/14/2004 | 09/15/2004 | 11/23/2005 | 2511556 | Premier Dental Products Co | foreign objects (polishing stones) in the tubes
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34782 | Z-1469-04 | Laryngectomy Tubes, Jackson Improved Stainless Steel, item number1037414 size 4, item number 1037415 size 5, item number 1037416 size 6, item number 1037417 size 7, item number 1037418 size 8, item number 1037419 size 9, and item number 1037420 size 10. | Jackson Improved Stainless Steel Laryngectomy Tubes | 2 | 09/14/2004 | 09/15/2004 | 11/23/2005 | 2511556 | Premier Dental Products Co | foreign objects (polishing stones) in the tubes
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34771 | Z-1458-04 | Tracheostomy Tubes, Jackson Original Stainless Steel, item number 1034300 size 00, item number 1034310 size 0, item number 1034311 size 1, item number 1034312 size 2, item number 1034313 size 3, item number 1034314 size 4, item number 1034315 size 5, item number 1034316 size 6, item number 1034317 size 7, item number 1034318 size 8, item number 1034319 size 9, and item number 1034320 size 10. | Jackson Original Stainless Steel Tracheostomy Tubes | 2 | 09/14/2004 | 09/15/2004 | 11/23/2005 | 2511556 | Premier Dental Products Co | foreign objects (polishing stones) in the tubes
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34779 | Z-1466-04 | Tracheostomy Tubes, Air Lon Inhalation Set (Inner Cannula with 15 mm Adapter) Nylon, item number 1050180 size 3, item number 1050182 size 4, item number 1050184 size 5, item number 1050186 size 6, item number 1050188 size 7, and item number 1050190 size 8. | Air Lon Inhalation Set (Inner Cannula with 15 mm Adapter) Nylon Tracheostomy Tubes | 2 | 09/14/2004 | 09/15/2004 | 11/23/2005 | 2511556 | Premier Dental Products Co | foreign objects (polishing stones) in the tubes
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34778 | Z-1465-04 | Tracheostomy Tubes, Air Lon Nylon, item number 1050150 size 3, item number 1050152 size 4, item number 1050154 size 5, item number 1050156 size 6, item number 1050158 size 7, and item number 1050160 size 8. | Air Lon Nylon Tracheostomy Tubes | 2 | 09/14/2004 | 09/15/2004 | 11/23/2005 | 2511556 | Premier Dental Products Co | foreign objects (polishing stones) in the tubes
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34777 | Z-1464-04 | Tracheostomy Tubes, Jackson Improved Extra Long Stainless Steel, item number 1034964 size 4, item number 1034965 size 5, item number 1034966 size 6, item number 1034967 size 7, item number 1034968 size 8, item number 1034969 size 9, and item number 1034970 size 10. | Jackson Improved Extra Long Stainless Steel Tracheostomy Tubes | 2 | 09/14/2004 | 09/15/2004 | 11/23/2005 | 2511556 | Premier Dental Products Co | foreign objects (polishing stones) in the tubes
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34776 | Z-1463-04 | Tracheostomy Tubes, Jackson Improved Short Stainless Steel with Adapter, item number 8100402 size 3,item number 8100040 size 4, item number 8100041 size 5, item number 8100229 size 6, item number 8100322 size 7, item number 8100374 size 8, item number 8100282 size 9, and item number 8100709 size10. | Jackson Improved Short Stainless Steel with Adapter Tracheostomy Tubes | 2 | 09/14/2004 | 09/15/2004 | 11/23/2005 | 2511556 | Premier Dental Products Co | foreign objects (polishing stones) in the tubes
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34773 | Z-1460-04 | Tracheostomy Tubes, Jackson Improved Stainless Steel, item number 1034900 size 00, item number 1034910 size 0, item number 1034911 size 1, item number 1034912 size 2, item number 1034913 size 3, item number 1034914 size 4, item number 1034915 size 5, item number 1034916 size 6, item number 1034917 size 7, item number 1034918 size 8, item number 1034919 size 9, and item number 1034920 size 10. | Jackson Improved Stainless SteelTracheostomy Tubes | 2 | 09/14/2004 | 09/15/2004 | 11/23/2005 | 2511556 | Premier Dental Products Co | foreign objects (polishing stones) in the tubes
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34774 | Z-1461-04 | Tracheostomy Tubes, Jackson Improved Stainless Steel with Adapter, item number 8100465 size 1, item number 8100373 size 3, item number 1034924 size 4, item number 1034925 size5, item number 1034926 size 6, item number 1034927 size 7, item number 1034928 size 8, item number 8100468 size 9, and item number 8100234 size 10. | Jackson Improved Stainless Steel with Adapter Tracheostomy Tubes | 2 | 09/14/2004 | 09/15/2004 | 11/23/2005 | 2511556 | Premier Dental Products Co | foreign objects (polishing stones) in the tubes
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34775 | Z-1462-04 | Tracheostomy Tubes, Jackson Improved Short Stainless Steel, item number 1034930 size 00, item number 1034940 size 0, item number 1034941 size 1, item number 1034942 size 2, item number 1034943 size 3, item number 1034944 size 4, item number 1034945 size 5, item number 1034946 size 6, item number 1034947 size 7, item number 1034948 size 8, item number 1034949 size 9, and item number 1034950 size 10. | Jackson Improved Short Stainless Steel Tracheostomy Tubes | 2 | 09/14/2004 | 09/15/2004 | 11/23/2005 | 2511556 | Premier Dental Products Co | foreign objects (polishing stones) in the tubes
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=34772 | Z-1459-04 | Tracheostomy Tubes, Jackson Original Short Stainless Steel, item number 1034600 size 00, item number 1034610 size 0, item number 1034611 size 1, item number 1034612 size 2, item number 1034613 size 3, item number 1034614 size 4,item number 1034615 size 5, item number 1034616 size 6, item number 1034617 size 7, item number 1034618 size 8, item number 1034619 size 9, and item number 1034620 size 10. | Jackson Original Short Stainless Steel Tracheostomy Tubes | 2 | 09/14/2004 | 09/15/2004 | 11/23/2005 | 2511556 | Premier Dental Products Co | foreign objects (polishing stones) in the tubes
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=183447 | Z-2946-2020 | Bivona Mid-Range Aire-Cuf Adult Tracheostomy Tube This tube is intended to provide direct airway access for a tracheotomized patient for up to 29 days. It may be reprocessed up to 10 times for single patient use. | Bivona Tracheostomy Tubes | 2 | 09/09/2020 |  | 05/13/2021 | 3012307300 | Smiths Medical ASD Inc. | Label has the incorrect size for the tracheotomy tube.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=111017 | Z-2174-2012 | Shiley Reusable Cannula Low Pressure Cuffed Tracheostomy Tubes, model numbers (Worldwide) 8LPC, 8FEN; (Outside U.S.) 8FENJ, 8FENJ-S, 8LPC-S, 8FEN-S. Product Usage: The product is intended for use to bypass upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. The product is sold as sterile. The packaging may indicate either Covidien or Tyco as the manufacturer as the firm has transitioned from Tyco to Covidien over the past 4 years. | Shiley | 1 | 08/10/2012 | 08/10/2012 | 01/22/2015 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is recalling certain cuffed Shiley tracheostomy tubes due to consumer complaints related to the connections and functions of the inner and outer cannula which were observed during mechanical ventilation of patients.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92611 | Z-2078-2010 | Cook Medical Ciaglia Blue Rhino G2 Advanced Percutaneous Tracheostomy Introducer Set, C-PTIS-100-HC-PERC6, G13165. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92656 | Z-2084-2010 | Cook Medical Ciaglia Blue Rhino G2 Advanced Percutaneous Tracheostomy Introducer Set, C-PTISY-100-HC-G-PERC6, G53178. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92648 | Z-2082-2010 | Cook Medical Ciaglia Blue Rhino G2 Advanced Percutaneous Tracheostomy Introducer Set, C-PTIS-100-HC-PERC8-CA, G36038. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92653 | Z-2083-2010 | Cook Medical Ciaglia Blue Rhino G2 Advanced Percutaneous Tracheostomy Introducer Set, C-PTISY-100-HC-PERC8, G12565. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92681 | Z-2090-2010 | Cook Medical Ciaglia Blue Dolphin Percutaneous Tracheostomy Introducer Tray, G36117, C-PTBS-2800-PERC8. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92610 | Z-2077-2010 | Cook Medical Ciaglia Blue Rhino G2 Advanced Percutaneous Tracheostomy Introducer Set, C-PTISY-100-A-HC-PERC8, G12566. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92626 | Z-2080-2010 | Cook Medical Ciaglia Blue Rhino G2 Advanced Percutaneous Tracheostomy Introducer Set, C-PTISY-100-HC-PERC6, G13166. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92676 | Z-2087-2010 | Cook Medical Ciaglia Blue Dolphin Percutaneous Tracheostomy Introducer Tray, G36116, C-PTBS-2600-PERC6. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92677 | Z-2088-2010 | Cook Medical Ciaglia Blue Dolphin Percutaneous Tracheostomy Introducer Tray, G36118, C-PTBSY-2600-PERC6. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92604 | Z-2076-2010 | Cook Medical Ciaglia Blue Rhino G2 Advanced Percutaneous Tracheostomy Introducer Set, C-PTIS-100-A-HC-PERC8, G12403. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92664 | Z-2086-2010 | Cook Medical Ciaglia Blue Rhino G2 Advanced Percutaneous Tracheostomy Introducer Set, C-PTIS-100-WCE-HC-G-PERC8, G53170. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92683 | Z-2092-2010 | Cook Medical Ciaglia Blue Dolphin Percutaneous Tracheostomy Introducer Tray, G51972, C-PTBSYJC-2800-PERC8. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92625 | Z-2079-2010 | Cook Medical Ciaglia Blue Rhino G2 Advanced Percutaneous Tracheostomy Introducer Set, C-PTIS-100-HC-PERC6-CA, G36037. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92627 | Z-2081-2010 | Cook Medical Ciaglia Blue Rhino G2 Advanced Percutaneous Tracheostomy Introducer Set, C-PTIS-100-HC-PERC8, G12523. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92680 | Z-2089-2010 | Cook Medical Ciaglia Blue Dolphin Percutaneous Tracheostomy Introducer Tray, G51971, C-PTBSYJC-2600-PERC6. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92663 | Z-2085-2010 | Cook Medical Ciaglia Blue Rhino G2 Advanced Percutaneous Tracheostomy Introducer Set, C-PTIS-100-WCE-HC-G-PERC6, G53169. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=92682 | Z-2091-2010 | Cook Medical Ciaglia Blue Dolphin Percutaneous Tracheostomy Introducer Tray, G49507, C-PTBSY-2800-PERC8. | Cook Medical Ciaglia Blue Dolphin | 1 | 07/26/2010 | 07/26/2010 | 01/23/2012 | 1820334 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=136460 | Z-1951-2015 | Shiley Neonatal, Pediatric and Long Pediatric Tracheostomy Tube with TaperGuard Cuff. Intended for use in providing tracheal access for airway managements. | Shiley Neonatal, Pediatric and Long Pediatric Tracheostomy Tube with TaperGuard Cuff | 1 | 07/09/2015 | 07/09/2015 | 04/12/2017 | 2936999 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Reports where patients who recently switched from the current Shiley Neonatal and Pediatric products to the affected products experienced discomfort and in some cases, breathing difficulties, corresponding to a negative effect on oxygen levels.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=136457 | Z-1950-2015 | Shiley Neonatal, Pediatric and Long Pediatric Tracheostomy Tube, Cuffless. Model Numbers: 2.5NEF,3.0NEF,3.5NEF,4.0NEF,4.5NEF,2.5PEF,3.0PEF,3.5PEF,4.0PEF,4.5PEF,5.0PEF,5.5PEF,5.0PELF,5.5PELF,6.0PELF,6.5PELF,2.5NCF,3.0NCF,3.5NCF,4.0NCF,4.5NCF,2.5PCF,3.0PCF,3.5PCF,4.0PCF,4.5PCF,5.0PCF,5.5PCF,5.0PLCF,5.5PLCF,6.0PLCF,6.5PLCF,3.0NEF-P,3.5NEF-P,4.0NEF-P,4.5NEF-P,3.0PEF-P,3.5PEF-P,4.0PEF-P,4.5PEF-P,5.0PEF-P,5.5PEF-P, | Shiley Neonatal, Pediatric and Long Pediatric Tracheostomy Tube, Cuffless | 1 | 07/09/2015 | 07/09/2015 | 04/12/2017 | 2936999 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Reports where patients who recently switched from the current Shiley Neonatal and Pediatric products to the affected products experienced discomfort and in some cases, breathing difficulties, corresponding to a negative effect on oxygen levels.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=187587 | Z-1993-2021 | PORTEX UNLTRperc Percutaneous Dilation Tracheostomy Kit, REF 100/596/090 - Product Usage: intended for use in a controlled setting such as an Intensive Care Unit or operating room with the assistance of trained personnel. | PORTEX UNLTRperc Percutaneous Dilation Tracheostomy Kit | 2 | 07/01/2021 |  | 06/09/2023 | 3012307300 | Smiths Medical ASD Inc. | One lot number of a specific model of ULTRAperc Percutaneous Dilation Tracheostomy Kit, 9.0mm may have been incorrectly labeled as a 7.0mm ID kit.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=181651 | Z-2485-2020 | Pediatric Adjustable Flange Tracheostomy Tube Cuffless | Bivona Cuffless Adjustable Neck Flange Hyperflex Pediatric Tracheostomy Tubes | 2 | 07/01/2020 |  | 05/05/2021 | 3012307300 | Smiths Medical ASD Inc. | One lot of 3.0mm Bivona Cuffless Adjustable Neck Flange Hyperflex Pediatric Tracheostomy Tubes was packaged with a 3.5mm ID adjustable neck flange plate.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=99651 | Z-2697-2011 | RUSCH Percutwist Set with Crystal Clear Sterile Crystal Clear Tracheostomy Tube is a device inserted into a patient's trachea surgically via the neck to maintain an open airway. | RUSCH Percutwist Set with Crystal Clear Sterile | 2 | 06/29/2011 | 06/29/2011 | 03/02/2012 | 3005747797 | Teleflex Medical | Recalled due to air leakage from the main ventilation lumen through air passageway, resulting in potential injury requiring medical intervention.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=99649 | Z-2695-2011 | RUSCH CrystalClear Tracheostomy Tube, Cuffed, Sterile Pack Crystal Clear Tracheostomy Tube is a device inserted into a patients trachea surgically via the neck to maintain an open airway. | RUSCH CrystalClear Tracheostomy Tube | 2 | 06/29/2011 | 06/29/2011 | 03/02/2012 | 3005747797 | Teleflex Medical | Recalled due to air leakage from the main ventilation lumen through air passageway, resulting in potential injury requiring medical intervention.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=99650 | Z-2696-2011 | RUSCH Crystal Clear PDT, Sterile Crystal Clear Tracheostomy Tube is a device inserted into a patients trachea surgically via the neck to maintain an open airway. | RUSCH Crystal Clear PDT | 2 | 06/29/2011 | 06/29/2011 | 03/02/2012 | 3005747797 | Teleflex Medical | Recalled due to air leakage from the main ventilation lumen through air passageway, resulting in potential injury requiring medical intervention.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=38630 | Z-0930-05 | Bivona brand Pediatric TTS Tracheostomy Tube, 3.5 mm tracheostomy tubes with a 3.5mm obturator ; product code 67P035. |  | 2 | 06/24/2005 | 06/25/2005 | 10/26/2005 | 1824231 | Smiths Medical ASD, Inc. | The obturator may be 3.0 mm in length instead of the labeled 3.5 mm.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=33190 | Z-1016-04 | Portex Inner Cannula for Tracheostomy Tube Size 7.0 mm, REF Code 566070 | Portex/Sims Portex | 2 | 06/14/2004 | 07/20/2004 | 05/02/2006 |  | Smiths Medical ASD, Inc | Incompatible products may cause potential trachea trauma
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=33193 | Z-1017-04 | Portex Inner Cannula for Tracheostomy Tube, 8.0 mm, REF Code 566080 | Portex/Sims Portex | 2 | 06/14/2004 | 07/20/2004 | 05/02/2006 |  | Smiths Medical ASD, Inc | Incompatible products may cause potential trachea trauma
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=39087 | Z-0868-05 | Portex Per-Fit Percutaneous Dilational Tracheostomy Kit, ID 8.0mm Product Number: 535080 |  | 2 | 06/03/2005 | 06/04/2005 | 05/08/2012 | 1217052 | Smiths Medical ASD, Inc. | Guiding catheter may have thinner and weaker sidewalls than designed, and may break.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=39091 | Z-0870-05 | Portex Per-Fit Percutaneous Dilational Tracheostomy Kit , ID 7.0mm Product Number: 538070E |  | 2 | 06/03/2005 | 06/04/2005 | 05/08/2012 | 1217052 | Smiths Medical ASD, Inc. | Guiding catheter may have thinner and weaker sidewalls than designed, and may break.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=39093 | Z-0872-05 | Portex Per-Fit Percutaneous Dilational Tracheostomy Kit , ID 9.0mm Product Number: 538090E |  | 2 | 06/03/2005 | 06/04/2005 | 05/08/2012 | 1217052 | Smiths Medical ASD, Inc. | Guiding catheter may have thinner and weaker sidewalls than designed, and may break.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=39089 | Z-0869-05 | Portex Per-Fit Percutaneous Dilational Tracheostomy Kit , ID 9.0mm Product Number: 535090 |  | 2 | 06/03/2005 | 06/04/2005 | 05/08/2012 | 1217052 | Smiths Medical ASD, Inc. | Guiding catheter may have thinner and weaker sidewalls than designed, and may break.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=39086 | Z-0867-05 | Portex Per-Fit Percutaneous Dilational Tracheostomy Kit ,ID 7.0mm Product Number: 535070 |  | 2 | 06/03/2005 | 06/04/2005 | 05/08/2012 | 1217052 | Smiths Medical ASD, Inc. | Guiding catheter may have thinner and weaker sidewalls than designed, and may break.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=39092 | Z-0871-05 | Portex Per-Fit Percutaneous Dilational Tracheostomy Kit , ID 8.0mm Product Number: 538080E |  | 2 | 06/03/2005 | 06/04/2005 | 05/08/2012 | 1217052 | Smiths Medical ASD, Inc. | Guiding catheter may have thinner and weaker sidewalls than designed, and may break.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=219873 | Z-2261-2026 | COVIDIEN Shiley Tracheostomy Tube Cuffed with Inner Cannula REF 6FEN UDI-DI code: 60884522007028 The devices are intended for use in providing tracheal access for airway management. Shiley Tracheostomy Tubes are double cannula tracheostomy tubes with reusable inner cannulae (LPC, FEN, CFS, CFN, LGT) and twist-lock connectors. These Shiley Tracheostomy Tubes have a radiopaque, biocompatible outer cannula constructed of polyvinyl chloride. | COVIDIEN Shiley Tracheostomy Tube Cuffed with Inner Cannula | 2 | 05/29/2026 | 05/29/2026 |  | 2936999 | Covidien LLC | Due to customer complaint regarding incorrect display box labeling.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=199621 | Z-1428-2023 | Blue Rhino G2-Multi Percutaneous Tracheostomy Introducer Tray, Reference Part Numbers C-PTISYJ-100-HC-G-EU-FLEX8.5 (G57725), C-PTISYJ-100-HC-G-NA-FLEX7.5 (G57720), C-PTISYJ-100-HC-G-NA-FLEX8.5 (G57721), C-PTISY-100-HC-G-NA-FLEX8.5 (G57717), C-PTISY-100-HC-G-NA-FLEX7.5 (G57716) | Blue Rhino G2-Multi Percutaneous Tracheostomy Introducer | 1 | 05/10/2023 | 05/10/2023 |  | 1820334 | Cook Incorporated | There were reports from customers that the tracheostomy in the sets and trays is not making a secure connection with the 15mm cap and other 15mm circuit components and accessories. This resulted in an unsecure connection between the device connector and circuit components, cap, or accessories.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=199620 | Z-1427-2023 | Blue Rhino G2-Multi Percutaneous Tracheostomy Introducer Set, Reference Part Numbers C-PTIS-100-HC-G-EU-FLEX7.5 (G57695), C-PTIS-100-HC-G-EU-FLEX8.5 (G57696), C-PTIS-100-HC-G-NA-FLEX7.5 (G57691), C-PTIS-100-HC-G-NA-FLEX8.5 (G57692) | Blue Rhino G2-Multi Percutaneous Tracheostomy Introducer | 1 | 05/10/2023 | 05/10/2023 |  | 1820334 | Cook Incorporated | There were reports from customers that the tracheostomy in the sets and trays is not making a secure connection with the 15mm cap and other 15mm circuit components and accessories. This resulted in an unsecure connection between the device connector and circuit components, cap, or accessories.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90624 | Z-1463-2010 | Shiley Disposable Cannula Low Pressure Cuffed Tracheostomy Tube, 6 DCT, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Disposable Cannula Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90627 | Z-1466-2010 | Shiley Fenestrated Low Pressure Cuffed Tracheostomy Tube, 4 FEN, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Fenestrated Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90630 | Z-1469-2010 | Shiley Fenestrated Low Pressure Cuffed Tracheostomy Tube, 10 FEN, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Fenestrated Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90634 | Z-1473-2010 | Shiley Disposable Cannula Fenestrated Low Pressure Cuffed Tracheostomy Tube, 10 DFEN, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Disposable Cannula Fenestrated Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90637 | Z-1475-2010 | Shiley Percutaneous Dual Cannula Tracheostomy Tube, 8 PERC, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Percutaneous Dual Cannula Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90638 | Z-1476-2010 | Shiley Fenestrated Low Pressure Cuffed Tracheostomy Tube with Evacuation Line, 6 FEN-S, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Fenestrated Low Pressure Cuffed Tracheostomy Tube with Evacuation Line | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90643 | Z-1481-2010 | Shiley Low Pressure Cuffed Tracheostomy Tube, 10 LPC, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90653 | Z-1489-2010 | Shiley Low Pressure Cuffed Tracheostomy Tube with Evacuation Line, 6 LPC-S, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Low Pressure Cuffed Tracheostomy Tube with Evacuation Line | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90880 | Z-1509-2010 | Shiley Specialized Tracheostomy Tube, M10SCT, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90667 | Z-1492-2010 | Shiley Cuffed Pediatric Tracheostomy Tube, 4.0 PDC, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Cuffed Pediatric Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90877 | Z-1506-2010 | Shiley Specialized Tracheostomy Tube, M7.5SCT, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90883 | Z-1512-2010 | Shiley Specialized Tracheostomy Tube, M5.5PDC, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90918 | Z-1521-2010 | Shiley Specialized Tracheostomy Tube, MDCT, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90921 | Z-1523-2010 | Shiley Specialized Tracheostomy Tube, M8DCT, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90642 | Z-1480-2010 | Shiley Low Pressure Cuffed Tracheostomy Tube, 8 LPC, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90672 | Z-1496-2010 | Shiley Cuffed Long Pediatric Tracheostomy Tube, 5.0 PLC, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Cuffed Pediatric Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90878 | Z-1507-2010 | Shiley Specialized Tracheostomy Tube, M8SCT, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90915 | Z-1519-2010 | Shiley Specialized Tracheostomy Tube, M6LPC, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90626 | Z-1465-2010 | Shiley Disposable Cannula Low Pressure Cuffed Tracheostomy Tube, 10 DCT, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Disposable Cannula Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90628 | Z-1467-2010 | Shiley Fenestrated Low Pressure Cuffed Tracheostomy Tube, 6 FEN, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Fenestrated Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90629 | Z-1468-2010 | Shiley Fenestrated Low Pressure Cuffed Tracheostomy Tube, 8 FEN, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Fenestrated Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90656 | Z-1490-2010 | Shiley Low Pressure Cuffed Tracheostomy Tube with Evacuation Line, 8 LPC-S, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Low Pressure Cuffed Tracheostomy Tube with Evacuation Line | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90657 | Z-1491-2010 | Shiley Low Pressure Cuffed Tracheostomy Tube with Evacuation Line, 10 LPC-S, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Low Pressure Cuffed Tracheostomy Tube with Evacuation Line | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90874 | Z-1503-2010 | Shiley Specialized Tracheostomy Tube, M6SCT, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90906 | Z-1513-2010 | Shiley Specialized Tracheostomy Tube, MPLC, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90916 | Z-1520-2010 | Shiley Specialized Tracheostomy Tube, M8LPC, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90632 | Z-1471-2010 | Shiley Disposable Cannula Fenestrated Low Pressure Cuffed Tracheostomy Tube, 6 DFEN, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Disposable Cannula Fenestrated Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90639 | Z-1477-2010 | Shiley Fenestrated Low Pressure Cuffed Tracheostomy Tube with Evacuation Line, 8 FEN-S, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90670 | Z-1494-2010 | Shiley Cuffed Pediatric Tracheostomy Tube, 5.0 PDC, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Cuffed Pediatric Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90673 | Z-1497-2010 | Shiley Cuffed Long Pediatric Tracheostomy Tube, 5.5 PLC, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Cuffed Pediatric Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90675 | Z-1498-2010 | Shiley Cuffed Long Pediatric Tracheostomy Tube, 6.0 PLC, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Cuffed Long Pediatric Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90882 | Z-1511-2010 | Shiley Specialized Tracheostomy Tube, M4.0PDC, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90911 | Z-1515-2010 | Shiley Specialized Tracheostomy Tube, MFEN, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90914 | Z-1518-2010 | Shiley Specialized Tracheostomy Tube, MLPC, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90625 | Z-1464-2010 | Shiley Disposable Cannula Low Pressure Cuffed Tracheostomy Tube, 8 DCT, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Disposable Cannula Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90631 | Z-1470-2010 | Shiley Disposable Cannula Fenestrated Low Pressure Cuffed Tracheostomy Tube, 4 DFEN, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Disposable Cannula Fenestrated Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90641 | Z-1479-2010 | Shiley Low Pressure Cuffed Tracheostomy Tube, 6 LPC, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90862 | Z-1501-2010 | Shiley Specialized Tracheostomy Tube, M5SCT, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90876 | Z-1505-2010 | Shiley Specialized Tracheostomy Tube, M7SCT, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90912 | Z-1516-2010 | Shiley Specialized Tracheostomy Tube, M6FEN, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90920 | Z-1522-2010 | Shiley Specialized Tracheostomy Tube, M6DCT, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90623 | Z-1462-2010 | Shiley Disposable Cannula Low Pressure Cuffed Tracheostomy Tube, 4 DCT, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Disposable Cannula Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90633 | Z-1472-2010 | Shiley Disposable Cannula Fenestrated Low Pressure Cuffed Tracheostomy Tube, 8 DFEN, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Disposable Cannula Fenestrated Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90651 | Z-1488-2010 | Shiley Low Pressure Cuffed Tracheostomy Tube with Evacuation Line, 4 LPC-S, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Low Pressure Cuffed Tracheostomy Tube with Evacuation Line | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90671 | Z-1495-2010 | Shiley Cuffed Pediatric Tracheostomy Tube, 5.5 PDC, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Cuffed Pediatric Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90861 | Z-1500-2010 | Shiley Specialized Tracheostomy Tube, MSCT, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90865 | Z-1502-2010 | Shiley Specialized Tracheostomy Tube, M5.5SCT, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90875 | Z-1504-2010 | Shiley Specialized Tracheostomy Tube, M6.5SCT, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90635 | Z-1474-2010 | Shiley Percutaneous Dual Cannula Tracheostomy Tube, 6 PERC, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Percutaneous Dual Cannula Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90640 | Z-1478-2010 | Shiley Low Pressure Cuffed Tracheostomy Tube, 4 LPC, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Low Pressure Cuffed Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90669 | Z-1493-2010 | Shiley Cuffed Pediatric Tracheostomy Tube, 4.5 PDC, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Cuffed Pediatric Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90679 | Z-1499-2010 | Shiley Cuffed Long Pediatric Tracheostomy Tube, 6.5 PLC, Assembled in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Cuffed Long Pediatric Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90879 | Z-1508-2010 | Shiley Specialized Tracheostomy Tube, M9SCT, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90881 | Z-1510-2010 | Shiley Specialized Tracheostomy Tube, MPDC, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90910 | Z-1514-2010 | Shiley Specialized Tracheostomy Tube, M5.5PLC, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=90913 | Z-1517-2010 | Shiley Specialized Tracheostomy Tube, M8FEN, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Intended for bypass of upper airway obstructions, provide long term ventilation, support and/or manage tracheal/bronchial secretions. | Shiley Specialized Tracheostomy Tube | 1 | 05/06/2010 | 05/06/2010 | 01/18/2012 | 2936999 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=199257 | Z-1357-2023 | Shiley Adult Flexible Tracheostomy Tube with TaperGuard Cuff, and Cuffless: Disposable Inner Cannula or Reusable Inner Cannula (REF, Product Description): 10CN10H, 10.0MM SHILEY CUFFED TRACH CAN; 10CN10R, 10.0MM ADT FLEX TRACH W TG CUFF; 10UN10A, 10.0MM UNCUFF TRACH TUBE X1; 10UN10H, 10.0MM SHILEY CUFFLESS TRACH C; 10UN10R, 10.0MM ADT FLEX TRACH CUFFLESS; 4CN65A, 6.5MM TRACH TUBE W TG CUFF X1; 4CN65H, 6.5MM SHILEY CUFFED TRACH CANN; 4CN65R, 6.5MM ADT FLEX TRACH W TG CUFF X1; 4UN65A, 6.5MM UNCUFF TRACH TUBE X1; 4UN65H, 6.5MM SHILEY CUFFLESS TRACH CA; 4UN65R, 6.5MM ADT FLEX TRACH CUFFLESSX1; 5CN70A, 7.0MM TRACH TUBE W TG CUFF X1; 5CN70H, 7.0MM SHILEY CUFFED TRACH CANN; 5CN70R, 7.0MM ADT FLEX TRACH W TG CUFF X1; 5UN70A, 7.0MM UNCUFF TRACH TUBE X1; 5UN70H, 7.0MM SHILEY CUFFLESS TRACH CA; 5UN70R, 7.0MM ADT FLEX TRACH CUFFLESSX1; 6CN75A, 7.5MM TRACH TUBE W TG CUFF X1; 6CN75H, 7.5MM SHILEY CUFFED TRACH CANN; 6CN75R, 7.5MM ADT FLEX TRACH W TG CUFF X1; 6UN75H, 7.5MM SHILEY CUFFLESS TRACH CA; 6UN75R, 7.5MM ADT FLEX TRACH CUFFLESSX1; 7CN80A, 8.0MM TRACH TUBE W TG CUFF X1; 7CN80H, 8.0MM SHILEY CUFFED TRACH CANN; 7CN80R, 8.0MM ADT FLEX TRACH W TG CUFF X1; 7UN80A, 8.0MM UNCUFF TRACH TUBE X1; 7UN80H, 8.0MM SHILEY CUFFLESS TRACH CA; 7UN80R, 8.0MM ADT FLEX TRACH CUFFLESSX1; 8CN85A, 8.5MM TRACH TUBE W TG CUFF X1; 8CN85H, 8.5MM SHILEY CUFFED TRACH CANN; 8CN85R, 8.5MM ADT FLEX TRACH W TG CUFF X1; 8UN85A, 8.5MM UNCUFF TRACH TUBE X1; 8UN85H, 8.5MM SHILEY CUFFLESS TRACH CA; 8UN85R, 8.5MM ADT FLEX TRACH CUFFLESSX1; 9CN90H, 9.0MM SHILEY CUFFED TRACH CANN; 9CN90R, 9.0MM ADT FLEX TRACH W TG CUFF X1; 9UN90H, 9.0MM SHILEY CUFFLESS TRACH CA; 9UN90R, 9.0MM ADT FLEX TRACH CUFFLESSX1 | Shiley" Adult Flexible Tracheostomy Tube with TaperGuard" Cuff and Cuffless | 1 | 04/19/2023 | 04/19/2023 |  | 2936999 | Covidien | A manufacturing error, resulted in a less than specified diameter of adult flexible tracheostomy tube connectors, which results in unsecure connections with 15mm caps and other 15mm circuit components and accessories. An unsecure connection could lead to respiratory failure, dyspnea, treatment delay, tissue injury, or bleeding.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=218921 | Z-1781-2026 | COOK MEDICAL Ciaglia Blue Rhino Percutaneous Tracheostomy Introducer Set: Reference Part Number C-PTIS-100-FLEX-HC, Order Number G32092; Reference Part Number C-PTIS-100-HC, Order Number G12115. | Ciaglia Blue Rhino Percutaneous Tracheostomy Introducer Set | 2 | 04/06/2026 | 04/06/2026 |  | 1820334 | Cook Incorporated | Products from the affected device lots were labeled with expiration dates that exceed the true shelf life.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=218920 | Z-1780-2026 | COOK MEDICAL Blue Rhino G2-Multi Percutaneous Tracheostomy Introducer Tray: Reference Part Number C-PTISY-100-HC-G-NA, Order Number G57703; Reference Part Number C-PTISY-100-UNL-HC-G-NA, Order Number G57706; Reference Part Number C-PTISY-100-UNS-HC-G-NA, Order Number G57704. | Blue Rhino G2-Multi Percutaneous Tracheostomy Introducer Tray | 2 | 04/06/2026 | 04/06/2026 |  | 1820334 | Cook Incorporated | Products from the affected device lots were labeled with expiration dates that exceed the true shelf life.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=218919 | Z-1779-2026 | COOK MEDICAL Blue Rhino G2-Multi Percutaneous Tracheostomy Introducer Set: Reference Part Number C-PTIS-100-HC-G-EU, Order Number G57685; Reference Part Number C-PTIS-100-HC-G-NA, Order Number G57682; Reference Part Number C-PTIS-100-UNL-HC-G-EU, Order Number G57687; Reference Part Number C-PTIS-100-UNS-HC-G-EU, Order Number G57686. | Blue Rhino G2-Multi Percutaneous Tracheostomy Introducer Set | 2 | 04/06/2026 | 04/06/2026 |  | 1820334 | Cook Incorporated | Products from the affected device lots were labeled with expiration dates that exceed the true shelf life.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=218922 | Z-1782-2026 | COOK MEDICAL Cook Staged Extubation Set: Reference Part Number C-CAE-14.0-83-EXTU, Order Number G24073 | Cook Staged Extubation Set | 2 | 04/06/2026 | 04/06/2026 |  | 1820334 | Cook Incorporated | Products from the affected device lots were labeled with expiration dates that exceed the true shelf life.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=212660 | Z-1390-2025 | Shiley Adult Flexible Tracheostomy Tube with TaperGuard Cuff Reusable Inner Cannula, REF: 7CN80R | Shiley Adult Flexible Tracheostomy Tube with TaperGuard Cuff Reusable Inner Cannula | 1 | 03/25/2025 | 03/25/2025 |  | 2936999 | Covidien | Tracheostomy tube with cuff reusable inner cannula has a flange that may disconnect from the outer cannula of the device, which could cause respiratory failure, aspiration, unspecified tissue injury, aspiration, respiratory tract infection, bronchospasm, a delay to treatment and/or death.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=106420 | Z-1255-2012 | Tracheotomy care kit, catalog# 900-110 | Tracheotomy care kit, | 2 | 03/21/2012 | 03/21/2012 | 04/12/2017 | 2648727 | Customed, Inc | The firm has determined that there is a possibility that packaging seals may be compromised. The issue could result in an injury to the patient due to product contamination or loss of sterility condition.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=153205 | Z-1465-2017 | Bivona TTS (Tight to Shaft) Pediatric Tracheostomy Tube, Model No. 67SP035, Reorder No. 67SP035. | Bivona TTS (Tight to Shaft) Pediatric Tracheostomy Tube | 2 | 03/14/2017 |  | 06/20/2018 | 3012307300 | Smiths Medical ASD Inc. | Smiths Medical became aware that one lot of the Bivona 3.5mm TTS Cuffed Pediatric tracheostomy tubes was incorrectly labeled for neonatal use instead of correctly labeled for pediatric use.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=78880 | Z-0997-2009 | Shiley PED3.0 Cuffless Pediatric Tracheostomy Tube, REF 3.0 PED, Sterile. Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Assembled in Mexico. The product is provided as one each shipped in an inner carton to end customers. The intended use of the device is to bypass upper airway obstructions, provide long term ventilation, and support and/or manage tracheal/bronchial secretions. | Shiley PED3.0 Cuffless Pediatric Tracheostomy Tube | 1 | 03/04/2009 | 03/04/2009 | 05/21/2012 | 2936999 | Covidien Limited | Difficulty inserting an obturator and/or suction catheter into the tracheostomy tube.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=30846 | Z-0569-04 | PER-FIT Percutaneous Dilational Tracheostomy Kit with Specialty Tracheostomy Tube REF # 511080 | SIMS Portex | 2 | 03/02/2004 | 07/20/2004 | 05/02/2006 |  | Smiths Medical ASD, Inc. | Lid tray incorrectly labeled as Tube I.D. 7.0 mm instead of correct size 8.0 mm
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=185515 | Z-1145-2021 | Tube Tracheostomy and Tube Cuff, Bivona Cuffless Neonatal/Pediatric FlexTend" Tracheostomy Tube (model 60PFS45, lots 3991171, 3991172, and 3988103) V Neck Flange 4.5mm - Product Usage: intended to provide direct airway access for a tracheotomized patient for up to 29 days. It may be reprocessed for single-patient use up to 5 times. | Bivona Cuffless Neonatal/Pediatric FlexTend" Tracheostomy Tube | 2 | 02/23/2021 |  | 12/20/2022 | 3012307300 | Smiths Medical ASD Inc. | Smiths Medical became aware that three lot numbers of a specific model of Bivona Cuffless FlexTend" Tracheostomy Tube (model 60PFS45, lots 3991171, 3991172, and 3988103) may have been incorrectly packaged with a straight neck flange instead of a V-neck flange. As a result, there is potential for patient injury if an incorrect device or accessory is selected for use. This use could lead to an internal injury resulting from excessive device contact with tissue, or an external patient injury due to external portions of the device pressing into or rubbing against tissue.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=161230 | Z-0577-2018 | Custom Made Tracheostomy Device The CADD Administration Sets with Flow Stop are designed to be used with a variety of CADD ambulatory infusion pumps to allow fluid delivery from an IV bag. | Custom Made Tracheostomy Device | 2 | 02/09/2018 |  | 04/24/2020 | 1824231 | Smiths Medical ASD, Inc. | Smiths Medical determined that a single custom Tracheostomy device had been shipped to the customer without required sterilization.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=43316 | Z-0395-06 | Bivona brand Mid-Range Aire-Cuf Extra Length Fixed Hyperflex Adult Tracheostomy Tube, I.D. 8.0 mm O.D. 11.0mm length 120 mm; REF 75FHXL80. | Hyperflex | 2 | 01/17/2006 | 01/18/2006 | 02/08/2007 | 1824231 | Smiths Medical ASD, Inc. | Sterility is compromised by incomplete package seal.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=43315 | Z-0394-06 | Bivona brand Mid-Range Aire-Cuf Adjustable Neck Flange Hyperflex Tracheostomy Tube, I.D. 6.0 mm O.D. 8.7 mm length 110 mm; REF 75HA60. | Hyperflex | 2 | 01/17/2006 | 01/18/2006 | 02/08/2007 | 1824231 | Smiths Medical ASD, Inc. | Sterility is compromised by incomplete package seal.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=43317 | Z-0396-06 | Bivona brand TTS Adjustable Neck Flange Hyperflex Tracheostomy Tube, I.D. 8.0 mm O.D. 11.0mm length 130 mm; REF 67HA80. | Hyperflex | 2 | 01/17/2006 | 01/18/2006 | 02/08/2007 | 1824231 | Smiths Medical ASD, Inc. | Sterility is compromised by incomplete package seal.
| JOH | JUBE TRACHEOSTOMY AND TUBE CUFF | 868.5800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfres/res.cfm?id=49177 | Z-0339-2007 | Bivona Adult Tracheostomy Tube, I.D. 8.5 mm, O.D 11.0 mm, Length 88 mm, TTS 11.8 mm, Mfrd by Smiths Medical Critical Care; Product code 670185. | Smiths Bivona Adult Trach Tube | 2 | 01/04/2007 | 01/05/2007 | 05/04/2007 | 1824231 | Smiths Medical ASD, Inc. | Mislabeled- The outer diameter (OD) of the shaft is labeled as 11.0 m on the inner, outer and tray label in error. The OD is actually 11.8 mm.